附件4

医疗器械生产质量管理规范 体外诊断试剂现场检查指导原则

| 章节 | 条款 | 内容 |
| --- | --- | --- |
| 机 构 和 人 员 | 1.1.1 | 应当建立与医疗器械生产相适应的管理机构，具备组织机构图。 查看提供的质量手册，是否包括企业的组织机构图，是否明确各部门的相互关系。 |
|  | *1.1.2 | 应当明确各部门的职责和权限，明确质量管理职能。 查看企业的质量手册，程序文件或相关文件，是否对各部门的职责权限作出规定；质量管理部门应当能独立行使职能, 查看质量管理部门的文件，是否明确规定对产品质量的相关事宜负有决策的权利。 |
|  | 1.1.3 | 生产管理部门和质量管理部门负责人不得互相兼任。 查看公司的任职文件或授权文件并对照相关生产、检验等履行职责的记录，核实是否与授权一致。 |
|  | 1.2.1 | 企业负责人应当是医疗器械产品质量的主要责任人。 |
|  | 1.2.2 | 企业负责人应当组织制定质量方针和质量目标。 查看质量方针和质量目标的制定程序、批准人员。 |
|  | 1.2.3 | 企业负责人应当确保质量管理体系有效运行所需的人力资源、基础设施和工作环境。 |
|  | 1.2.4 | 企业负责人应当组织实施管理评审，定期对质量管理体系运行情况进行评估，并持续改进。 查看管理评审文件和记录，核实企业负责人是否组织实施管理评审。 |
|  | *1.2.5 | 企业负责人应当确保企业按照法律、法规和规章的要求组织生产。 |
|  | 1.3.1 | 企业负责人应当确定一名管理者代表。 查看管理者代表的任命文件。 |
|  | *1.3.2 | 管理者代表应当负责建立、实施并保持质量管理体系，报告质量管理体系的运行情况和改进需求，提高员工满足法规、规章和顾客要求的意识。 查看是否对上述职责作出明确规定。查看管理者代表报告质量管理体系运行情况和改进的相关记录 |
|  | 1.4.1 | 技术、生产、质量管理部门负责人应当熟悉医疗器械法律法规，具有质量管理的实践经验，应当有能力对生产管理和质量管理中实际问题作出正确判断和处理。 查看相关部门负责人的任职资格要求，是否对专业知识、工作技能、工作经历作出规定；查看考核评价记录，现场询问，确定是否符合要求。 |
|  | 1.5.1 | 应当配备与生产产品相适应的专业技术人员、管理人员和操作人员。 查看相关人员的资格要求。 |
|  | *1.5.2 | 应当具有相应的质量检验机构或专职检验人员。 查看组织机构图、部门职责要求、岗位人员任命等文件确认是否符合要求。 |
|  | *1.6.1 | 从事影响产品质量工作的人员，应当经过与其岗位要求相适应的培训，具有相关的理论知识和实际操作技能。 应当确定影响医疗器械质量的岗位，规定这些岗位人员所必须具备的专业知识水平（包括学历要求）、工作技能、工作经验。查看培训内容、培训记录和考核记录，是否符合要求。 |
|  | 1.7.1 | 体外诊断试剂生产、技术和质量管理人员应当具有医学、检验学、生物学、免疫学或药学等与所生产产品相关的专业知识，并具有相应的实践经验，以确保具备在生产、质量管理中履行职责的能力。 查看企业对相关岗位人员的任职要求、学历证书或培训等材料，是否符合要求。 |
|  | 1.8.1 | 凡在洁净室（区）工作的人员应当定期进行卫生和微生物学基础知识、洁净作业等方面培训。 查看培训计划和记录,是否对在洁净室（区）工作的人员定期进行了卫生和微生物学基础知识、洁净技术等方面的培训。 |
|  | 1.8.2 | 临时进入洁净室（区）的人员，应当对其进行指导和监督。 查看是否制定了相关文件，对临时进入洁净室的人员 (包括外来人员)进出洁净区的指导和监督作出了规定。 |
|  | 1.9.1 | 从事体外诊断试剂生产的全体人员，包括清洁、维修等人员均应当根据其产品和所从事的生产操作进行专业和安全防护培训。 查看培训计划和记录,是否对从事体外诊断试剂生产的全体人员，包括清洁、维修等人员根据其产品和所从事的生产操作进行了专业和安全防护培训。 企业从事高生物活性、高毒性、强传染性、强致敏性等有特殊要求产品的生产和质量检验人员应当具备相关岗位操作资格或接受相关专业技术培训和防护知识培训，合格后方可上岗。 |
|  | 1.10.1 | 应当建立对人员的清洁要求，制定洁净室（区）工作人员卫生守则。 查看工作人员卫生守则,是否对人员清洁、进出程序、洁净服的穿戴作出规定。 |
|  | 1.10.2 | 人员进入洁净室（区）应当按照程序进行净化，并穿戴工作帽、口罩、洁净工作服、工作鞋。 现场观察人员进入洁净室（区）是否按照程序进行净化，并按规定正确穿戴工作帽、口罩、洁净工作服、工作鞋或鞋套。 |
|  | 1.10.3 | 裸手接触产品的操作人员每隔一定时间应当对手再次进行消毒。裸手消毒剂的种类应当定期更换。 查看消毒剂配制或领用记录，是否按要求定期更换裸手消毒剂的种类。 |
|  | 1.11.1 | 应当制定人员健康要求，建立人员健康档案。 查看人员健康要求的文件，是否对人员健康要求作出规定，并建立人员健康档案。 |
|  | 1.11.2 | 直接接触物料和产品的操作人员每年至少体检一次。患有传染性和感染性疾病的人员不得从事直接接触产品的工作。 查看洁净间直接接触物料和产品的人员的体检报告或健康证明,是否按规定时间进行体检，患有传染性和感染性疾病的人员未从事直接接触产品的工作。 |
|  | 1.12.1 | 应当明确人员服装要求，制定洁净和无菌工作服的管理规定。工作服及其质量应当与生产操作的要求及操作区的洁净度级别相适应，其式样和穿着方式应当能够满足保护产品和人员的要求。无菌工作服应当能够包盖全部头发、胡须及脚部，并能阻留人体脱落物。 查看洁净和无菌工作服的管理规定；现场观察服装的符合性及人员穿戴的符合性。 |
|  | 1.12.2 | 洁净工作服和无菌工作服不得脱落纤维和颗粒性物质。 查看洁净工作服和无菌工作服是否选择质地光滑、不易产生静电、不脱落纤维和颗粒性物质的材料制作。 |
| 厂房与设施 | 2.1.1 | 厂房与设施应当符合产品的生产要求。 |
|  | 2.1.2 | 生产、行政和辅助区的总体布局应当合理，不得互相妨碍。 |
|  | *2.2.1 | 厂房与设施应当根据所生产产品的特性、工艺流程及相应的洁净级别要求进行合理设计、布局和使用。 |
|  | 2.2.2 | 生产环境应当整洁、符合产品质量需要及相关技术标准的要求。 |
|  | 2.2.3 | 产品有特殊要求的，应当确保厂房的外部环境不能对产品质量产生影响，必要时应当进行验证。 |
|  | 2.3.1 | 厂房应当确保生产和贮存产品质量以及相关设备性能不会直接或间接地受到影响。 |
|  | 2.3.2 | 厂房应当有适当的照明、温度、湿度和通风控制条件。 |
|  | 2.4.1 | 厂房与设施的设计和安装应当根据产品特性采取必要措施，有效防止昆虫或其他动物进入。 现场查看是否配备了相关设施。 |
|  | 2.4.2 | 对厂房与设施的维护和维修不应影响产品质量。 |
|  | 2.5.1 | 生产区应当有足够空间，并与产品生产规模、品种相适应。 |
|  | 2.6.1 | 仓储区应当能够满足原材料、包装材料、中间品、产品等贮存条件和要求。 |
|  | 2.6.2 | 仓储区应当按照待验、合格、不合格、退货或召回等进行有序、分区存放各类材料和产品，便于检查和监控。 现场查看是否设置了相关区域并进行了标识，对各类物料是否按规定区域存放，应当有各类物品的贮存记录。 |
|  | *2.7.1 | 应当配备与产品生产规模、品种、检验要求相适应的检验场所和设施。 对照产品生产工艺的要求和产品检验要求以及检验方法，核实企业是否具备相关检测条件。 |
|  | 2.8.1 | 应当有整洁的生产环境。厂区的地面、路面周围环境及运输等不应对产品的生产造成污染。厂区应当远离有污染的空气和水等污染源的区域。 现场查看生产环境，应当整洁、无积水和杂草。厂区的地面、路面周围环境及运输等不应对产品的生产造成污染。检查地面、道路平整情况及减少露土、扬尘的措施和厂区的绿化，以及垃圾、闲置物品等的存放情况。 |
|  | 2.8.2 | 行政区、生活区和辅助区的总体布局合理，不得对生产区有不良影响。 |
|  | 2.9.1 | 生产厂房应当设置防尘、防止昆虫和其他动物进入的设施。洁净室（区）的门、窗及安全门应当密闭，洁净室（区）的门应当向洁净度高的方向开启。 |
|  | 2.10.1 | 应当根据体外诊断试剂的生产过程控制，确定在相应级别的洁净室（区）内进行生产的过程，避免生产中的污染。 查看相关文件，是否明确了生产过程的洁净度级别；现场查看是否在相应级别洁净室（区）内进行生产，是否能避免生产中的污染。 |
|  | 2.10.2 | 空气洁净级别不同的洁净室（区）之间的静压差应当大于5帕，洁净室（区）与室外大气的静压差应大于10帕，并应当有指示压差的装置。 现场查看是否配备了指示压差的装置，空气洁净级别不同的洁净室（区）之间以及洁净室（区）与室外大气的静压差是否符合要求。 |
|  | 2.10.3 | 相同级别洁净室间的压差梯度应当合理。 查看相关文件，是否明确相同级别洁净室间的压差梯度，现场查看是否符合要求。 |
|  | 2.11.1 | 酶联免疫吸附试验试剂、免疫荧光试剂、免疫发光试剂、聚合酶链反应（PCR）试剂、金标试剂、干化学法试剂、细胞培养基、校准品与质控品、酶类、抗原、抗体和其他活性类组分的配制及分装等产品的配液、包被、分装、点膜、干燥、切割、贴膜以及内包装等，生产区域应当不低于100,000级洁净度级别。 |
|  | 2.12.1 | 阴性或阳性血清、质粒或血液制品等的处理操作，生产区域应当不低于10，000级洁净度级别，并应当与相邻区域保持相对负压。 |
|  | 2.13.1 | 无菌物料等分装处理操作，操作区域应当符合局部100级洁净度级别。 |
|  | 2.14.1 | 普通类化学试剂的生产应当在清洁环境中进行。 |
|  | 2.15.1 | 洁净室（区）空气洁净度级别应当符合下表规定: 洁净室（区）空气洁净度级别表 洁净度级 别; 尘粒最大允许数／m3; 微生物最大允许数. ≥0.5μm; ≥5μm; 浮游菌／m3; 沉降菌／皿. 100级; 3,500; 0; 5; l. 10,000级; 350,0O0; 2,000; 100; 3. 1OO,000级; 3,500,000; 20,000; 500; 10. |
|  | 2.16.1 | 洁净室（区）应当按照体外诊断试剂的生产工艺流程及所要求的空气洁净度级别进行合理布局，人流、物流走向应当合理。同一洁净室（区）内或相邻洁净室（区）间的生产操作不得互相交叉污染。 现场查看洁净室（区）的人流、物流走向是否合理，是否能够避免交叉污染。 |
|  | 2.17.1 | 进入洁净室（区）的管道、进回风口布局应当合理，水、电、气输送线路与墙体接口处应当可靠密封，照明灯具不得悬吊。 |
|  | 2.18.1 | 洁净室（区）的温度和相对湿度应当与产品生产工艺要求相适应。无特殊要求时，温度应当控制在18～28℃，相对湿度控制在45%～65%。 现场查看温湿度装置及记录，是否符合要求。 |
|  | 2.19.1 | 洁净室（区）和非洁净室（区）之间应有缓冲设施。 |
|  | 2.20.1 | 洁净室（区）的内表面（墙面、地面、天棚、操作台等）应当平整光滑、无裂缝、接口严密、无颗粒物脱落，避免积尘，并便于清洁处理和消毒。 |
|  | 2.21.1 | 洁净室（区）的空气如循环使用应当采取有效措施避免污染和交叉污染。 查看空调净化系统送、回、排风相关图纸，空气如循环使用的，核实控制措施是否有效。 |
|  | 2.22.1 | 洁净室（区）内的水池、地漏应安装防止倒灌的装置，避免对环境和物料造成污染。 |
|  | 2.22.2 | 100级的洁净室（区）内不得设置地漏。 |
|  | 2.23.1 | 产尘操作间应当保持相对负压或采取有效措施，防止粉尘扩散，避免交叉污染。 |
|  | *2.24.1 | 对具有污染性、传染性和高生物活性的物料应当在受控条件下进行处理，避免造成传染、污染或泄漏等。 |
|  | *2.25.1 | 生产激素类、操作有致病性病原体或芽胞菌制品的，应当使用单独的空气净化系统，与相邻区域保持负压，排出的空气不能循环使用。 现场查看生产激素类、操作有致病性病原体、芽胞菌制品的区域与其他类产品生产区域是否严格分开。如系多楼层建筑，在同一生产层面与其他一般品种的生产线不能共用物料通道、人员通道、包装线等，防止产品交叉污染。 现场查看是否为独立的空气净化系统，排出的空气不能循环使用。 现场查看压差表，是否与相邻区域保持负压。 |
|  | 2.26.1 | 进行危险度二级及以上的病原体操作应当配备生物安全柜，空气应当进行过滤处理后方可排出。 |
|  | 2.26.2 | 应当对过滤器的性能进行定期检查以保证其有效性。 查看维护保养记录，是否进行定期检查。 |
|  | 2.26.3 | 使用病原体类检测试剂的阳性血清应当有相应的防护措施。 查看相关文件，是否明确了防护措施。 现场查看阳性血清的保存条件及使用记录。 |
|  | *2.27.1 | 对于特殊的高致病性病原体的采集、制备，应当按照有关部门颁布的行业标准如人间传染病微生物名录、微生物和生物医学实验室生物安全通用准则、实验室生物安全通用要求等相关规定，配备相应的生物安全设施。 查看相应文件，是否明确病原体的危险等级，生物安全防护要求是否符合相关规定。 现场查看是否配备生物安全设施，是否符合防护要求。 |
|  | *2.28.1 | 生产聚合酶链反应（PCR）试剂的，其生产和检验应当在独立的建筑物或空间内，保证空气不直接联通，防止扩增时形成的气溶胶造成交叉污染。生产聚合酶链反应（PCR）试剂的，其生产和质检的器具不得混用，用后应严格清洗和消毒。 |
|  | 2.29.1 | 洁净室（区）内的人数应当与洁净室（区）面积相适应。 查看验证记录，是否对现场工作人员数量上限进行验证，确认能够满足洁净控制要求。核实现场工作人员数量，查看相关记录，不应超过验证时所确认的现场工作人员数量上限。 |
|  | 2.30.1 | 对生产环境没有空气净化要求的体外诊断试剂，应当在清洁环境内进行生产。 清洁条件的基本要求：要有防尘、通风、防止昆虫或其他动物以及异物混入等措施；人流、物流分开，人员进入生产车间前应当有换鞋、更衣、佩戴口罩和帽子、洗手、手消毒等清洁措施；生产场地的地面应当便于清洁，墙、顶部应平整、光滑，无颗粒物脱落；操作台应当光滑、平整、无缝隙、耐腐蚀，便于清洗、消毒；应当对生产区域进行定期清洁、清洗和消毒；应当根据生产要求对生产车间的温湿度进行控制。 |
|  | *2.31.1 | 易燃、易爆、有毒、有害、具有污染性或传染性、具有生物活性或来源于生物体的物料的管理应当符合国家相关规定。所涉及的物料应当列出清单，专区存放、专人保管和发放，并制定相应的防护规程。 查看清单，是否将易燃、易爆、有毒、有害、具有污染性或传染性、具有生物活性或来源于生物体的物料均已纳入。 查看防护规程，是否明确管理、防护要求。 现场查看是否专区、专人管理，是否有发放记录。 |
|  | 2.32.1 | 动物室应当在隔离良好的建筑体内，与生产、质检区分开，不得对生产造成污染。 |
| 设 备 | *3.1.1 | 应当配备与所生产产品和规模相匹配的生产设备、工艺装备，应当确保有效运行。 对照生产工艺流程图，查看设备清单，所列设备是否满足生产需要；核查现场设备是否与设备清单相关内容一致；应当制定设备管理制度。 |
|  | 3.2.1 | 生产设备的设计、选型、安装、维修和维护应当符合预定用途，便于操作、清洁和维护。 查看生产设备验证记录，确认是否满足预定要求。现场查看生产设备是否便于操作、清洁和维护。 |
|  | 3.2.2 | 生产设备应当有明显的状态标识，防止非预期使用。 现场查看生产设备标识。 |
|  | 3.2.3 | 应当建立生产设备使用、清洁、维护和维修的操作规程，并保存相应的设备操作记录。 |
|  | *3.3.1 | 应当配备与产品检验要求相适应的检验仪器和设备，主要检验仪器和设备应当具有明确的操作规程。 对照产品检验要求和检验方法，核实企业是否具备相关检测设备。主要检测设备是否制定的了操作规程。 |
|  | 3.4.1 | 应当建立检验仪器和设备的使用记录，记录内容应当包括使用、校准、维护和维修等情况。 |
|  | 3.5.1 | 应当配备适当的计量器具，计量器具的量程和精度应当满足使用要求，计量器具应当标明其校准有效期，保存相应记录。 查看计量器具的校准记录，确定是否在有效期内使用。 |
|  | 3.6.1 | 洁净室（区）空气净化系统应当经过确认并保持连续运行，维持相应的洁净度级别，并在一定周期后进行再确认。 查看洁净室（区）空气净化系统的确认和再确认记录。 |
|  | 3.6.2 | 若停机后再次开启空气净化系统，应当进行必要的测试或验证，以确认仍能达到规定的洁净度级别要求。 如果洁净室（区）空气净化系统不连续使用，应当通过验证明确洁净室（区）空气净化系统重新启用的要求，并查看每次启用空气净化系统前的操作记录是否符合控制要求。 如果未进行验证，在停机后再次开始生产前应当对洁净室（区）的环境参数进行检测，确认达到相关标准要求。 |
|  | 3.7.1 | 应当确定所需要的工艺用水，当生产过程中使用工艺用水时，应当配备相应的制水设备，并有防止污染的措施，用量较大时应当通过管道输送至洁净室（区）的用水点。工艺用水应当满足产品质量的要求。 现场查看制水设备是否满足用水要求；是否具有防止污染的措施。 |
|  | 3.8.1 | 应当制定工艺用水的管理文件。 查看工艺用水的管理文件，是否明确工艺用水种类是否符合《药典》或GB/T6682或YY/T1244等标准要求。是否包括设备维护、保养、清洗、消毒，水质监测、检测的要求。 |
|  | 3.8.2 | 工艺用水的储罐和输送管道应当满足所生产的产品对于水质的要求，并定期清洗、消毒。 现场查看工艺用水的储罐和输送管道应当用不锈钢或其他无毒材料制成，应当定期清洗、消毒并进行记录。 |
|  | 3.9.1 | 配料罐容器与设备连接的主要固定管道应当标明内存的物料名称、流向，定期清洗和维护，并标明设备运行状态。 现场查看主要固定管道标识是否清晰、是否包含物料名称和流向，与其连接的设备是否标识了设备运行状态。 查看主要固定管道的清洗和维护记录。 |
|  | 3.10.1 | 与物料或产品直接接触的设备、容器具及管道表面应当光洁、平整、无颗粒物质脱落、无毒、耐腐蚀，不与物料或产品发生化学反应和粘连，易于清洁处理和消毒或灭菌。 |
|  | 3.11.1 | 需要冷藏、冷冻的原料、半成品、成品，应当配备相应的冷藏、冷冻储存设备，并按规定监测设备运行状况、记录储存温度。 查看相关文件，是否明确需要冷藏、冷冻的原料、半成品和成品的储存条件。 现场查看是否配备冷藏、冷冻的设施设备，如有冷库，查看冷库是否有温度显示、是否有报警功能、如冷库断电是否有应急措施。 查看储存环境的温度记录，是否持续满足产品的储存要求。 |
|  | 3.11.2 | 冷藏、冷冻体外诊断试剂应当配备符合其温度要求的运输设施设备; 查看冷藏、冷冻体外诊断试剂运输管理文件，是否明确试剂运输过程中的防护要求和方法。 查看现场，是否具有相应的运输设施设备，是否符合储存运输的温度要求。 |
| 文 件 管理 | *4.1.1 | 应当建立健全质量管理体系文件，包括质量方针和质量目标、质量手册、程序文件、技术文件和记录，以及法规要求的其他文件。 质量方针应当在企业内部得到沟通和理解；应当在持续适宜性方面得到评审。质量目标应当与质量方针保持一致；应当根据总的质量目标，在相关职能和层次上进行分解，建立各职能和层次的质量目标；应当包括满足产品要求所需的内容；应当可测量、可评估；应当有具体的方法和程序来保障。 |
|  | 4.1.2 | 质量手册应当对质量管理体系作出规定。 查看企业的质量手册，应当包括企业质量目标、组织机构及职责、质量体系的适用范围和要求。 |
|  | 4.1.3 | 程序文件应当根据产品生产和质量管理过程中需要建立的各种工作程序而制定，包含本规范所规定的各项程序文件。 |
|  | *4.1.4 | 技术文件应当包括产品技术要求及相关标准、生产工艺规程、作业指导书、检验和试验操作规程、安装和服务操作规程等相关文件。 |
|  | 4.2.1 | 应当建立文件控制程序，系统地设计、制定、审核、批准和发放质量管理体系文件。 |
|  | 4.2.2 | 文件的起草、修订、审核、批准、替换或撤销、复制、保管和销毁等应当按照控制程序管理，并有相应的文件分发、撤销、复制和销毁记录。 |
|  | 4.2.3 | 文件更新或修订时应当按规定评审和批准，能够识别文件的更改和修订状态。 查看相关记录确认文件的更新或修订是否经过评审和批准；其更改和修订状态是否能够得到识别。 |
|  | 4.2.4 | 分发和使用的文件应当为适宜的文本，已撤销或作废的文件应当进行标识，防止误用。 到工作现场抽查现场使用的文件，确认是否是有效版本。作废文件是否明确标识。 |
|  | 4.3.1 | 应当确定作废的技术文件等必要的质量管理体系文件的保存期限，满足产品维修和产品质量责任追溯等需要。 保存期限应当不少于企业所规定的医疗器械寿命期。 |
|  | 4.4.1 | 应当建立记录控制程序，包括记录的标识、保管、检索、保存期限和处置要求等。 |
|  | 4.4.2 | 记录应当保证产品生产、质量控制等活动可追溯性。 |
|  | 4.4.3 | 记录应当清晰、完整，易于识别和检索，防止破损和丢失。 |
|  | 4.4.4 | 记录不得随意涂改或销毁，更改记录应当签注姓名和日期，并使原有信息仍清晰可辨，必要时，应当说明更改的理由。 |
|  | 4.4.5 | 记录的保存期限至少相当于生产企业所规定的医疗器械的寿命期，但从放行产品的日期起不少于2年，或符合相关法规要求，并可追溯。 |
| 设计开发 | 5.1.1 | 应当建立设计控制程序并形成文件，对医疗器械的设计和开发过程实施策划和控制。 查看设计控制程序文件，应当清晰、可操作，能控制设计开发过程，至少包括以下内容： 1.设计和开发的各个阶段的划分； 2.适合于每个设计和开发阶段的评审、验证、确认和设计转换活动; 3.设计和开发各阶段人员和部门的职责、权限和沟通； 4.风险管理要求。 |
|  | 5.2.1 | 在进行设计和开发策划时，应当确定设计和开发的阶段及对各阶段的评审、验证、确认和设计转换等活动，应当识别和确定各个部门设计和开发的活动和接口，明确职责和分工。 查看设计和开发策划资料，应当根据产品的特点，对设计开发活动进行策划，并将策划结果形成文件。至少包括以下内容： 1.设计和开发项目的目标和意义的描述，技术指标分析； 2.确定了设计和开发各阶段，以及适合于每个设计和开发阶段的评审、验证、确认和设计转换活动； 3．应当识别和确定各个部门设计和开发的活动和接口，明确各阶段的人员或组织的职责、评审人员的组成，以及各阶段预期的输出结果； 4.主要任务和阶段性任务的策划安排与整个项目的一致； 5.确定产品技术要求的制定、验证、确认和生产活动所需的测量装置； 6.风险管理活动。 应当按照策划实施设计和开发，当偏离计划而需要修改计划时，应当对计划重新评审和批准。 |
|  | 5.3.1 | 设计和开发输入应当包括预期用途规定的功能、性能和安全要求、法规要求、风险管理控制措施和其他要求。 |
|  | 5.3.2 | 应当对设计和开发输入进行评审并得到批准，保持相关记录。 |
|  | *5.4.1 | 设计和开发输出应当满足输入要求，包括采购、生产和服务所需的相关信息、产品技术要求等。 查看设计和开发输出资料，至少符合以下要求： 1.采购信息，如原材料、包装材料、组件和部件技术要求； 2.生产和服务所需的信息，如产品图纸（包括零部件图纸）、工艺配方、作业指导书、环境要求等； 3.产品技术要求； 4.产品检验规程或指导书； 5.规定产品的安全和正常使用所必须的产品特性，如产品使用说明书、包装和标签要求等。产品使用说明书是否与注册申报和批准的一致； 6.标识和可追溯性要求； 7.提交给注册审批部门的文件，如研究资料、产品技术要求、注册检验报告、临床评价资料（如有）、医疗器械安全有效基本要求清单等； 8.样机或样品； 9.生物学评价结果和记录，包括材料的主要性能要求。 |
|  | 5.4.2 | 设计和开发输出应当得到批准，保持相关记录。 |
|  | 5.5.1 | 应当在设计和开发过程中开展设计和开发到生产的转换活动，以使设计和开发的输出在成为最终产品规范前得以验证，确保设计和开发输出适用于生产。 查看相关文件，至少符合以下要求： 1.应当在设计和开发过程中开展设计转换活动以解决可生产性、部件及材料的可获得性、所需的生产设备、操作人员的培训等； 2.设计转换活动应当将产品的每一技术要求正确转化成与产品实现相关的具体过程或程序； 3.设计转换活动的记录应当表明设计和开发输出在成为最终产品规范前得到验证，并保留验证记录，以确保设计和开发的输出适于生产； 4. 应当对特殊过程的转换进行确认，确保其结果适用于生产，并保留确认记录。 |
|  | 5.6.1 | 应当在设计和开发的适宜阶段安排评审，保持评审结果及任何必要措施的记录。 查看相关文件和记录，至少符合以下要求： 1.应当按设计开发策划的结果，在适宜的阶段进行设计和开发评审； 2.应当保持设计和开发评审记录，包括评审结果和评审所采取必要措施的记录。 |
|  | 5.7.1 | 应当对设计和开发进行验证，以确保设计和开发输出满足输入的要求，并保持验证结果和任何必要措施的记录。 查看相关文件和记录，至少符合以下要求： 1.应当结合策划的结果，在适宜的阶段进行设计和开发验证，确保设计开发输出满足输入的要求； 2.应当保持设计和开发验证记录、验证结果和任何必要措施的记录； 3.若设计和开发验证采用的是可供选择的计算方法或经证实的设计进行比较的方法，应当评审所用的方法的适宜性，确认方法是否科学和有效。 |
|  | 5.8.1 | 应当对设计和开发进行确认，以确保产品满足规定的使用要求或者预期用途的要求，并保持确认结果和任何必要措施的记录。 查看相关文件和记录，至少符合以下要求： 1.应当在适宜阶段进行设计和开发确认，确保产品满足规定的使用要求或预期用途的要求； 2.设计和开发确认活动应当在产品交付和实施之前进行； 3.应当保持设计和开发确认记录，包括临床评价或临床试验的记录，保持确认结果和任何必要措施的记录。 |
|  | 5.9.1 | 确认可采用临床评价或者性能评价。进行临床试验时应当符合医疗器械临床试验法规的要求。 查看临床评价报告及其支持材料。若开展临床试验的，其临床试验应当符合法规要求并提供相应的证明材料。对于需要进行临床评价或性能评价的医疗器械，应当能够提供评价报告和（或）材料。 |
|  | 5.10.1 | 应当对设计和开发的更改进行识别并保持记录。 |
|  | 5.10.2 | 必要时，应当对设计和开发更改进行评审、验证和确认，并在实施前得到批准。 查看设计和开发更改的评审记录，至少符合以下要求： 1.应当包括更改对产品组成部分和已交付产品的影响； 2.设计和开发更改的实施应符合医疗器械产品注册的有关规定； 3.设计更改的内容和结果涉及到改变医疗器械产品注册证（备案凭证）所载明的内容时，企业应当进行风险分析，并按照相关法规的规定，申请变更注册（备案），以满足法规的要求。 |
|  | *5.10.3 | 当选用的材料、零件或者产品功能的改变可能影响到医疗器械产品安全性、有效性时，应当评价因改动可能带来的风险，必要时采取措施将风险降低到可接受水平，同时应当符合相关法规的要求。 |
|  | 5.11.1 | 应当在包括设计和开发在内的产品实现全过程中，制定风险管理的要求并形成文件，保持相关记录。 查看风险管理文件和记录，至少符合以下要求： 1.风险管理应当覆盖企业开发的产品实现的全过程。 2.应当建立对医疗器械进行风险管理的文件，保持相关记录，以确定实施的证据。 3.应当将医疗器械产品的风险控制在可接受水平。 |
|  | 5.12.1 | 研制条件，包括配合使用的设备、仪器和试剂应当满足研究所需，研制所用的设备、仪器和试剂应当保存使用记录。 |
|  | 5.12.2 | 研制过程中主要原料、中间体、重要辅料应当明确来源，其数量、使用量及其剩余量应当保存记录。 |
|  | 5.12.3 | 工艺研究、技术要求/分析性能研究、稳定性研究、检验、临床试验/评价（包括预实验）研究、参考值研究等各个阶段的样品数量、贮存条件、留样、使用或销毁情况应当保存记录，样品试制量应当满足从事研究所需要的数量。 |
| 采 购 | *6.1.1 | 应当建立采购控制程序。 采购程序内容至少包括：采购流程、合格供应商的选择、评价和再评价规定、采购物品检验或验证的要求、采购记录的要求。 |
|  | *6.1.2 | 应当确保采购物品符合规定的要求，且不低于法律法规的相关规定和国家强制性标准的相关要求。 |
|  | 6.2.1 | 应当根据采购物品对产品的影响，确定对采购物品实行控制的方式和程度。 查看对采购物品实施控制方式和程度的规定，核实控制方式和程度能够满足产品要求。 |
|  | 6.3.1 | 应当建立供应商审核制度，对供应商进行审核评价。必要时，应当进行现场审核。 是否符合《医疗器械生产企业供应商审核指南》的要求。 |
|  | 6.3.2 | 应当保留供方评价的结果和评价过程的记录。 |
|  | *6.4.1 | 应当与主要原材料供应商签订质量协议，明确双方所承担的质量责任。 |
|  | 6.5.1 | 采购时应当明确采购信息，清晰表述采购要求，包括采购物品类别、验收准则、规格型号、规程、图样等内容。 从采购清单中抽查相关采购物品的采购要求，确认是否符合本条要求。 |
|  | 6.5.2 | 应当建立采购记录，包括采购合同、原材料清单、供应商资质证明文件、质量标准、检验报告及验收标准等。 |
|  | *6.5.3 | 采购记录应当满足可追溯要求。 |
|  | 6.6.1 | 应当对采购物品进行检验或验证，确保满足生产要求。 查看采购物品的检验或验证记录。 |
|  | *6.7.1 | 外购的标准品、校准品、质控品、生产用或质控用血液的采购应满足可追溯要求。 |
|  | *6.7.2 | 应当由企业或提供机构测定病原微生物及明确定值范围；应当对其来源地、定值范围、灭活状态、数量、保存、使用状态等信息有明确记录，并由专人负责。 查看相应文件，是否明确记录的要求和专人负责的要求。 查看病原微生物的相关记录，是否符合要求。 |
| 生 产 管 理 | *7.1.1 | 应当按照建立的质量管理体系进行生产，以保证产品符合强制性标准和经注册或者备案的产品技术要求。 |
|  | *7.2.1 | 应当编制生产工艺规程、作业指导书等，明确关键工序和特殊过程。 查看相关文件；是否明确关键工序和特殊过程，对关键工序和特殊过程的重要参数是否做验证或确认的规定。 |
|  | 7.3.1 | 在生产过程中需要对原材料、中间品等进行清洁处理的，应当明确清洁方法和要求，并对清洁效果进行验证。 |
|  | 7.4.1 | 应当根据生产工艺特点对环境进行监测，并保存记录。 |
|  | 7.5.1 | 应当对生产的特殊过程进行确认，并保存记录，包括确认方案，确认方法、操作人员、结果评价、再确认等内容。 |
|  | 7.5.2 | 生产过程中采用的计算机软件对产品质量有影响的，应当进行验证或确认。 |
|  | *7.6.1 | 每批（台）产品均应当有生产记录，并满足可追溯的要求。 |
|  | 7.6.2 | 生产记录应当包括：产品名称、规格型号、原材料批号、生产批号或产品编号、生产日期、数量、主要设备、工艺参数、操作人员等内容。 |
|  | 7.7.1 | 应当建立产品标识控制程序，用适宜的方法对产品进行标识，以便识别，防止混用和错用。 |
|  | *7.8.1 | 应当在生产过程中标识产品的检验状态，防止不合格中间产品流向下道工序。 查看是否对检验状态标识方法作出规定，现场查看生产过程中的检验状态标识，是否符合文件规定。 |
|  | *7.9.1 | 应当建立产品的可追溯性程序，规定产品追溯范围、程度、标识和必要的记录。 |
|  | *7.10.1 | 产品的说明书、标签应当符合相关法律法规及标准要求。 |
|  | 7.11.1 | 应当建立产品防护程序，规定产品及其组成部分的防护要求，包括污染防护、静电防护、粉尘防护、腐蚀防护、运输防护等要求。防护应当包括标识、搬运、包装、贮存和保护等。 现场查看产品防护程序是否符合规范要求；现场查看并抽查相关记录，确认产品防护符合要求。 |
|  | 7.12.1 | 洁净室（区）内使用的压缩空气等工艺用气均应当经过净化处理。 现场查看工艺用气的净化处理装置及管路设置,工艺用气是否经过净化处理。 |
|  | 7.12.2 | 与产品使用表面直接接触的气体，其对产品的影响程度应当进行验证和控制，以适应所生产产品的要求。 查看对与产品使用表面直接接触的气体对产品所造成的影响进行评价和验证的记录，是否根据评价和验证的结果规定了控制措施并实施。 |
|  | 7.13.1 | 生产设备、容器具等应当符合洁净环境控制和工艺文件的要求。 查看生产设备和容器具的相关文件，是否明确其净化要求，符合环境控制和工艺文件的要求。 |
|  | 7.14.1 | 应当按照物料的性状和储存要求进行分类存放管理，应当明确规定中间品的储存条件和期限。 查看相关文件，是否对物料进行分类，明确分类存放的要求和中间品储存条件、期限。 现场查看物料存放是否分类存放，满足其储存要求。 现场查看中间品的储存条件和期限是否符合规定。 |
|  | 7.14.2 | 物料应当在规定的使用期限内、按照先进先出的原则使用，无规定使用期限的，应当根据物料的稳定性数据确定储存期限。 查看物料管理相关文件，是否明确先进先出的使用原则，是否对无规定使用期限的物料明确根据物料的稳定性数据确定储存期限。 查看物料领用记录，是否符合先进先出的要求。对无规定使用期限的物料，查看物料的稳定性数据。 |
|  | 7.14.3 | 储存期内发现存储条件变化且可能影响产品质量时，应及时进行复验。 查看物料储存环境记录，对不符合储存要求的，查看复验记录。 |
|  | 7.15.1 | 进入洁净室（区）的物品应当按程序进行净化处理。 查看相关文件，是否明确物品进入洁净室（区）的相关程序。 现场查看进入洁净室（区）物品的净化程序是否符合规定要求。 |
|  | 7.16.1 | 在生产过程中，应当建立产品标识和生产状态标识控制程序。 |
|  | 7.16.2 | 对现场各类物料和生产区域、设备、管路的状态进行识别和管理。 |
|  | 7.17.1 | 应当对每批产品中关键物料进行物料平衡核查。如有显著差异，必须查明原因，在得出合理解释，确认无潜在质量事故后，方可按正常产品处理。 |
|  | 7.18.1 | 应当制定批号管理制度，对主要物料、中间品和成品按规定进行批号管理，并保存和提供可追溯的记录。 查看批号管理制度，是否明确主要物料、中间品和成品的批号编制方法。 查看相关记录，是否可追溯到主要物料、中间品。 |
|  | 7.18.2 | 同一试剂盒内各组分批号不同时应尽量将生产日期接近的组分进行组合，应当在每个组分的容器上均标明各自的批号和有效期。 |
|  | 7.18.3 | 整个试剂盒的有效期应以最先到有效期的组分的效期为准。 |
|  | 7.19.1 | 不同品种产品的生产应当做到有效隔离，以避免相互混淆和污染。 |
|  | 7.19.2 | 有数条包装线同时进行包装时，应当采取隔离或其他有效防止混淆的措施。 |
|  | 7.20.1 | 应当制定洁净室（区）的卫生管理文件。按照规定对洁净室（区）进行清洁处理和消毒，并作好记录。 查看洁净室（区）工艺卫生管理文件和记录，工艺卫生管理文件是否包含下列内容：1.设备清洁规定；2.工位器具（容器具、管路、储罐等）清洁规定；3.物料清洁规定；4.操作台、场地、墙壁、顶棚清洁规定；5.清洁工具的清洁及存放规定；6.洁净室（区）空气消毒规定；7.消毒剂选择、使用的管理规定。 现场查看洁净室（区）内的清洁卫生工具，是否使用无脱落物、易清洗、易消毒，是否按用途分类使用，不同洁净室（区）的清洁工具不得跨区使用。 |
|  | 7.20.2 | 所用的消毒剂或消毒方法不得对设备、容器具、物料和产品造成污染。 查看相关文件，是否对消毒剂或消毒方法作出规定,应包括消毒剂品种、使用时间、频次、更换周期等内容，应保留相关的记录。 应当对所选择的消毒方法、选用的消毒剂进行效果评价或验证；所用的消毒剂或消毒方法不应当对设备、容器具、物料和产品造成污染。 |
|  | 7.20.3 | 消毒剂品种应当定期更换，防止产生耐药菌株。 现场查看所使用的消毒剂是否符合文件规定，是否按期进行更换。 |
|  | 7.21.1 | 生产设备所用的润滑剂、清洗剂均不得对产品造成污染。 查看验证报告，所使用的润滑剂、清洗剂等不会对产品造成污染，或有相应措施消除污染的影响。 |
|  | 7.22.1 | 应当建立清场的管理规定。 是否根据生产工艺制定清场的管理规定及记录。 |
|  | 7.22.2 | 前一道工艺结束后或前一种产品生产结束后必须进行清场，确认合格后才可以入场进行其他生产，并保存清场记录。 现场查看上次生产遗留物的处理是否符合规定，查看清场记录。 |
|  | 7.22.3 | 相关的配制和分装器具必须专用，使用后进行清洗、干燥等洁净处理。 |
|  | 7.23.1 | 应当建立可追溯性程序并形成文件，应当规定可追溯性的范围、程度、标识和记录。 |
|  | 7.23.2 | 记录应当包括生产过程所用的原材料、生产过程、生产设备、操作人员和生产环境等内容。 |
|  | 7.24.1 | 生产一定周期后，应当对关键项目进行再验证。 查看验证相关文件，是否明确关键项目再验证的要求，查看相关记录。 |
|  | 7.24.2 | 当影响产品质量的主要因素，如工艺、质量控制方法、主要原辅料、主要生产设备等需要开展重新验证的条件发生改变时，应当进行相关内容的重新验证。 查看验证相关文件，是否明确上述因素发生改变时再验证的要求，查看相关记录。 |
|  | 7.24.3 | 企业应当根据不同产品特性提出验证的时间。 |
|  | 7.25.1 | 生产车间停产超过十二个月，重新组织生产前应当对生产环境及设施设备、主要原辅材料、关键工序、检验设备及质量控制方法等重新进行验证。 查看相关文件和记录，是否符合要求。 |
|  | 7.25.2 | 不足十二个月时，如有必要，企业也应当重新对生产环境和设施设备进行验证。 |
|  | 7.26.1 | 应当对生产用需要灭活的血清或血浆建立灭活处理的操作规程。 查看灭活处理的操作规程，是否明确生产用需要灭活的血清或血浆的种类，灭活的方法，操作过程中的防护措施、灭活前后物料的隔离和标识要求。 查看灭活记录。 |
|  | 7.26.2 | 按照操作规程的要求，对生产用灭活前后的血清或血浆状态进行明显的区分和标识。 现场查看是否对灭活前后的血清或血浆进行区分和标识，标识应有血清或血浆的名称、批号、是否灭活、灭活日期等。 |
|  | 7.27.1 | 生产中的废液、废物等应当进行无害化处理，并符合相关的环保要求。 |
| 质量控制 | 8.1.1 | 应当建立质量控制程序，规定产品检验部门、人员、操作等要求。 查看质量控制程序，是否对产品的检验部门职责、人员资质、检验操作规程等作出规定。 |
|  | 8.1.2 | 应当规定检验仪器和设备的使用、校准等要求，以及产品放行的程序。 查看质量控制程序，是否对检验仪器、设备的使用和校准作出规定。 |
|  | 8.2.1 | 应当定期对检验仪器和设备进行校准或检定，并予以标识。 查看检验仪器和设备是否按规定实施了校准或检定，是否进行了标识。 |
|  | 8.2.2 | 应当规定检验仪器和设备在搬运、维护、贮存期间的防护要求，防止检验结果失准。 |
|  | 8.2.3 | 当发现检验仪器和设备不符合要求时，应当对以往检验结果进行评价，并保存验证记录。 查看设备使用、维护记录，当检验仪器设备不符合要求的情况，是否对以往检测的结果进行了评价，并保存相关记录。 |
|  | 8.2.4 | 对用于检验的计算机软件，应当进行确认。 |
|  | *8.3.1 | 应当根据强制性标准以及经注册或者备案的产品技术要求制定产品的检验规程，并出具相应的检验报告或证书。 查看产品检验规程是否涵盖强制性标准以及经注册或者备案的产品技术要求的性能指标；确认检验记录是否能够证实产品符合要求；查看是否根据检验规程及检验结果出具相应的检验报告或证书。 |
|  | 8.3.2 | 需要常规控制的进货检验、过程检验和成品检验项目原则上不得进行委托检验。对于检验条件和设备要求较高，确需委托检验的项目，可委托具有资质的机构进行检验，以证明产品符合强制性标准和经注册或者备案的产品技术要求。 |
|  | *8.4.1 | 每批（台）产品均应当有批检验记录，并满足可追溯要求。 |
|  | 8.4.2 | 检验记录应当包括进货检验、过程检验和成品检验的检验记录、检验报告或证书等。 |
|  | *8.5.1 | 应当规定产品放行程序、条件和放行批准要求。 查看产品放行程序，是否明确了放行的条件和放行批准的要求。应当规定有权放行产品人员及其职责权限，并应当保持批准的记录。 |
|  | 8.5.2 | 放行的产品应当附有合格证明。 |
|  | 8.6.1 | 应当根据产品和工艺特点制定留样管理规定，按规定进行留样，并保持留样观察记录。 |
|  | 8.7.1 | 应当建立校准品、参考品量值溯源程序。对每批生产的校准品、参考品进行赋值。 查看校准品、参考品量值溯源程序（参见ISO17511）和校准品、参考品的赋值记录。 |
|  | *8.8.1 | 生产和检验用的菌毒种应当标明来源，验收、储存、保管、使用、销毁应执行国家有关医学微生物菌种保管的规定和病原微生物实验室生物安全管理条例。应当建立生产用菌毒种的原始种子批、主代种子批和工作种子批系统。 |
|  | *8.9.1 | 生产用细胞应当建立原始细胞库、主代细胞库、工作细胞库。应当建立细胞库档案资料和细胞操作日志。 |
|  | *8.9.2 | 自行制备抗原或抗体，应当对所用原料的来源和性质有详细的记录并可追溯。 |
|  | 8.10.1 | 应当对检验过程中使用的标准品、校准品、质控品建立台帐及使用记录。应当记录其来源、批号、效期、溯源途径、主要技术指标、保存状态等信息。 |
|  | 8.10.2 | 按照规定进行标准品、校准品、质控品复验并保存记录。 |
|  | *8.11.1 | 留样应当在规定条件下储存。应当建立留样台帐，及时记录留样检验信息，留样检验报告应当注明留样批号、效期、检验日期、检验人、检验结果等。 查看是否制定了留样管理制度，是否明确了留样的数量、规格、、储存条件、检验要求等。 现场查看留样室（区）的环境是否满足留样产品的要求，是否配备所需的环境监测设备是否有记录。 查看留样台账和留样检验报告，是否符合要求。 |
|  | 8.11.2 | 留样期满后应当对留样检验报告进行汇总、分析并归档。 |
| 销售 和售后 服 务 | *9.1.1 | 应当建立产品销售记录，并满足可追溯要求。 |
|  | 9.1.2 | 销售记录至少应当包括：医疗器械名称、规格、型号、数量、生产批号、有效期、销售日期、购货单位名称、地址、联系方式等内容。 |
|  | 9.2.1 | 直接销售自产产品或者选择医疗器械经营企业，应当符合医疗器械相关法规和规范要求。 |
|  | 9.2.2 | 发现医疗器械经营企业存在违法违规经营行为时，应当及时向当地食品药品监督管理部门报告。 |
|  | 9.3.1 | 应当具备与所生产产品相适应的售后服务能力，建立健全售后服务制度。 |
|  | 9.3.2 | 应当规定售后服务要求并建立售后服务记录，并满足可追溯的要求。 |
|  | 9.4.1 | 需要由企业安装的医疗器械，应当确定安装要求和安装验证的接收标准，建立安装和验收记录。 |
|  | 9.4.2 | 由使用单位或其他企业进行安装、维修的，应当提供安装要求、标准和维修零部件、资料、密码等，并进行指导。 |
|  | 9.5.1 | 应当建立顾客反馈处理程序，对顾客反馈信息进行跟踪分析。 查看程序文件是否对上述活动的实施作出了规定，并对顾客反馈信息进行了跟踪和分析。 |
| 不 合 格 品 控 制 | 10.1.1 | 应当建立不合格品控制程序，规定不合格品控制的部门和人员的职责与权限。 |
|  | *10.2.1 | 应当对不合格品进行标识、记录、隔离、评审，根据评审结果，应当对不合格品采取相应的处置措施。 现场查看不合格品的标识、隔离是否符合程序文件的规定，抽查不合格品处理记录，是否按文件的规定进行评审。 |
|  | 10.3.1 | 在产品销售后发现产品不合格时，应及时采取相应措施，如召回、销毁等措施。现场查看在产品销售后发现不合格时的处置措施，是否召回和销毁等。 |
|  | 10.4.1 | 不合格品可以返工的，企业应当编制返工控制文件。返工控制文件应当包括作业指导书、重新检验和重新验证等内容。 查看返工控制文件，是否对可以返工的不合格品作出规定；抽查返工活动记录，确认是否符合返工控制文件的要求。 |
|  | 10.4.2 | 不能返工的，应当建立相关处置制度。 |
| 不良事件监测 、 分 析 和 改 进 | 11.1.1 | 应当指定相关部门负责接收、调查、评价和处理顾客投诉，并保持相关记录。 查看有关职责权限的文件，确定是否对上述活动作出规定。 |
|  | *11.2.1 | 应当按照有关法规要求建立医疗器械不良事件监测制度，开展不良事件监测和再评价工作，保持相关记录。 查看企业建立的不良事件的监测制度，是否规定了可疑不良事件管理人员的职责、报告原则、上报程序、上报时限，制定了启动实施医疗器械再评价的程序和文件等，并符合法规要求。查看相关记录，确认是否存在不良事件，并按规定要求实施。 |
|  | 11.3.1 | 应当建立数据分析程序，收集分析与产品质量、不良事件、顾客反馈和质量管理体系运行有关的数据，验证产品安全性和有效性，并保持相关记录。 查看数据分析的实施记录，是否按程序规定进行，是否应用了统计技术并保留了数据分析结果的记录。 |
|  | 11.4.1 | 应当建立纠正措施程序，确定产生问题的原因，采取有效措施，防止相关问题再次发生。 |
|  | 11.4.2 | 应当建立预防措施程序，确定潜在问题的原因，采取有效措施，防止问题发生。 |
|  | *11.5.1 | 对存在安全隐患的医疗器械，应当按照有关法规要求采取召回等措施，并按规定向有关部门报告。 |
|  | 11.6.1 | 应当建立产品信息告知程序，及时将产品变动、使用等补充信息通知使用单位、相关企业或消费者。 |
|  | 11.7.1 | 应当建立质量管理体系内部审核程序，规定审核的准则、范围、频次、参加人员、方法、记录要求、纠正预防措施有效性的评定等内容，以确保质量管理体系符合本规范的要求。 查看内部审核程序是否包括了上述内容。查看内审资料，实施内审的人员是否经过培训，内审的记录是否符合要求，针对内审发现的问题是否采取了纠正措施，是否有效。 |
|  | *11.8.1 | 应当定期开展管理评审，对质量管理体系进行评价和审核，以确保其持续的适宜性、充分性和有效性。 查看管理评审文件和记录，应包括管理评审计划、管理评审报告以及相关改进措施，管理评审报告中是否包括了对法规符合性的评价。是否在规定时间内进行了管理评审，是否提出了改进措施并落实具体职责和要求，是否按计划实施。 |

注解：

1.本指导原则条款编号的编排方式为：X1.X2.X3，其中X1为章节的顺序号，如1.1.1的第一位X1表示“机构与人员”章节，2.1.1的第一位X1表示“厂房与设施”章节；X2为同一章节内条款的顺序号，如1.1.1的第二位X2表示“机构与人员”章节第一条要求，1.2.1的第二位X2表示“机构与人员”章节第二条要求；X3为同一条款内细化的检查指导的顺序号，如1.1.1的第三位X3表示“机构与人员”章节对第一条要求细化的第一个检查要点，1.1.2的第三位X3表示“机构与人员”章节对第一条要求细化的第二个检查要点。

其他章节编号规则相同。

2.每一章节的检查指导原则由前后两部分组成，每章的前半部分是按照《医疗器械生产质量规范》所规定条款制定的检查指导，每章的后半部是按照《医疗器械生产质量规范附录体外诊断试剂》所规定条款制定的检查指导。

国家食品药品监督管理总局办公厅      2015年9月25日印发

Annex 4

Good Manufacturing Practice for Medical Devices

Guiding Principles for On-Site Inspection of In Vitro Diagnostic Reagents

| Chapter | Terms | Content |
| --- | --- | --- |
| Organizations and Personnel | 1.1.1 | Administrative organizations with the organization chart which are adapted to medical device manufacture shall be established. View the quality manual to see whether the organization chart of the enterprise is included, and whether the relationship between departments is clear. |
|  | *1.1.2 | The responsibilities and authorities of each department, and the quality management functions shall be specified. Check the quality manual, procedure documents or related documents of the enterprise, to see whether the rights and responsibilities of various departments are provided; the quality management department shall be able to function independently; check documents of the quality management department, to see whether decision-making authority is provided for matters related to product quality. |
|  | 1.1.3 | Heads of production management department and quality management department shall not hold mutually concurrent posts. Check appointment or authorization documents of the enterprise by taking reference to the duty execution records related to production and inspection to verify whether they are consistent with the authorization. |
|  | 1.2.1 | Head of the manufacturer shall be the main responsible persons for medical device product quality. |
|  | 1.2.2 | Head of the manufacturer shall organize the development of quality policy and quality objectives. Check the development procedures and approvers of the quality policy and quality objectives |
|  | 1.2.3 | Head of the manufacturer shall ensure the required human resource, infrastructure and work environment for effective operation of the quality management system. |
|  | 1.2.4 | Head of the manufacturer shall organize the implementation of the management review, the assessment of quality management system on a regular basis, and the continuous improvement. Check the management review documents and records, to verify whether head of the manufacturer have organized the implementation of management reviews. |
|  | *1.2.5 | Head of the manufacturer shall ensure that the enterprise has organized production in accordance with the requirements of laws, regulations and rules. |
|  | 1.3.1 | Head of the manufacturer shall designate a management representative. Check the appointment documents of the management representative. |
|  | *1.3.2 | The management representative shall be responsible for establishing, implementing and maintaining quality management systems, reporting the operation and improvement requirements of the quality management system, improving employee awareness of regulations and rules compliance, and customer requirements. Check to see whether the above responsibilities have been clearly defined. Check relevant reports of the management representative on operations and improvements of the quality management. |
|  | 1.4.1 | Heads of the technical, production and quality management departments shall be familiar with the laws and regulations of medical devices, with practical experience in quality management, and shall be capable of making correct judgment and dealing with real issues in production management and quality management. Check the qualification requirements of relevant heads of departments, to see whether requirements on professional knowledge, skills, and work experience are provided; check assessment records, carry out on-the-spot inquiry to determine whether the requirements are met. |
|  | 1.5.1 | Professional and technical personnel, management personnel and operating personnel fitting in with production of the product shall be equipped with. Check qualification requirements of relevant personnel. |
|  | *1.5.2 | Appropriate quality inspection organizations or full-time inspectors shall be equipped with. Check the organization chart, the departmental responsibility requirements, staff appointments and other documents to confirm whether the requirements are met. |
|  | *1.6.1 | Personnel performing work affecting product quality shall be subject to appropriate training about their job requirements, and shall have relevant theoretical knowledge and practical skills. Identify the jobs affecting the medical device quality, specify the level of expertise (including academic requirements), job skills and work experience for personnel holding these jobs. Check training content, training and assessment records, to see whether the requirements are met. |
|  | 1.7.1 | In vitro diagnostic reagent production, technology and quality management personnel shall have the knowledge of medicine, laboratory science, biology, immunology, or pharmacy and other professional knowledge of production of the product, and have the appropriate practical experience, in order to ensure the ability to fulfill responsibilities of production and quality management. Check appointment requirements, qualifications or training materials of post holders of the enterprise to see whether requirements are met. |
|  | 1.8.1 | Training on basic knowledge of hygiene and microbiology, and clean operation shall be carried out for personnel working in clean rooms (areas) on a regular basis. Check training schedules and records, to see whether personnel working in clean rooms (areas) have accepted regular training on basic knowledge of hygiene and microbiology, and clean operation. |
|  | 1.8.2 | Personnel temporarily entering the clean rooms (areas) shall be guided and supervised. Check to see whether relevant guidance and supervision for personnel (including visitors) temporarily entering the clean room have been provided. |
|  | 1.9.1 | All personnel engaged in production of in vitro diagnostic reagents, including cleaning and maintenance personnel shall accept professional and safety trainings in terms of their products and production operations. Check training schedules and records, to see whether professional and safety trainings in terms of their products and the production operations have been carried out for all those who are engaged in the production of in vitro diagnostic reagents, including cleaning and maintenance personnel. Production and quality inspection personnel of the enterprise engaged in production of products of highly active, toxic, infectious or sensitizing, and products with special requirements shall have relevant post operation qualifications, or accept relevant professional and technical training and protection knowledge training, and shall be qualified before taking the jobs. |
|  | 1.10.1 | Cleaning requirements shall be established for personnel, and health code shall be developed for personnel working in the clean room (area). Check staff health code, to see whether cleanliness, access, and clean clothes of personnel are specified. |
|  | 1.10.2 | Personnel entering the clean room (area) shall be purified in accordance with procedures, and shall wear hats, masks, clean garments, and shoes. Inspect on-site to see whether personnel entering the clean room (area) are purified in accordance with procedures, and are appropriately wearing hats, masks, clean garments, work boots or shoes covers in accordance with the regulations. |
|  | 1.10.3 | Operators contacting products with bare-hands shall disinfect their hands at regular intervals. Types of bare hand disinfectants shall be altered on a regular basis. Check disinfectant preparation or use records, to see whether the types of bare hand disinfectants are altered on a regular basis. |
|  | 1.11.1 | Staff health requirements shall be developed, and staff health records shall be established. Check the documents on staff health requirements, to see whether provisions for staff health requirements are made, and staff health records are established. |
|  | 1.11.2 | Operator in direct contact with materials and products shall accept medical examinations for at least once every year. Personnel suffering from contagious and infectious diseases shall not engage in jobs in direct contact with the product. Check the medical examination reports or health certificates of personnel in direct contact with products in the clean room, to see whether they have accepted medical examinations on a regular basis, and ensure personnel suffering from contagious and infectious diseases are not engaged in jobs in direct contact with the product. |
|  | 1.12.1 | Personnel clothing requirements shall be specified, and clean and aseptic garment regulations shall be developed. The garment and its quality shall be appropriate to the production requirements and the cleanliness level of the operating areas, its style and wearing way shall be able to meet the requirements of protection of products and personnel. The aseptic garments shall be able to cover completely hair, beard and feet and prevent body materials from being left. Check regulations for clean and aseptic garments; inspect on-site the clothing compliance and compliance of personnel wearing of clothing. |
|  | 1.12.2 | Clean garments and aseptic garments must not shed fibers and particulate matters. Check to see whether materials of smooth texture, which are difficult to produce static electricity, and do not shed fibers and particulate matters are selected for making clean garments and aseptic garments |
| Premises and Facilities | 2.1.1 | Premises and facilities shall meet the production requirements. |
|  | 2.1.2 | The general layout of the production, administrative and support areas shall be reasonable and must not hinder each other. |
|  | *2.2.1 | Facilities and equipment shall be designed, laid out and used appropriately in accordance with the properties of the manufactured product and its process. |
|  | 2.2.2 | The production environment shall be clean, consistent with product quality needs and requirements of the relevant technical standards. |
|  | 2.2.3 | For products with special requirements, it shall be ensured that the building's external environment does not have an impact on product quality and, if necessary, shall be verified. |
|  | 2.3.1 | It shall be ensured that, premises will not directly or indirectly affect the quality of products being produced or stored, and the performance of relevant facilities. |
|  | 2.3.2 | Premises shall have proper lighting, temperature, humidity and ventilation control conditions. |
|  | 2.4.1 | Necessary measures shall be taken based on the design and installation of premises and facilities to effectively afford maximum protection against the entry of insects or other animals. Inspect on-site to see whether related facilities are equipped with. |
|  | 2.4.2 | Maintenance and repair of premises and facilities shall not affect product quality. |
|  | 2.5.1 | The production area shall have adequate space and shall be compatible with the production scale, and product varieties. |
|  | 2.6.1 | The storage area shall be able to meet the storage conditions and requirements of raw materials, packaging materials, intermediates, and products. |
|  | 2.6.2 | Orderly partition storage of various materials and products shall be carried out based on the classifications of in quarantine, released and rejected, returned or recalled, to facilitate inspection and supervision. Inspect on-site to see whether relevant areas have been set up and labeled, and whether various materials are stored in the required storage areas. Storage records of all kinds of products shall be kept. |
|  | *2.7.1 | Inspection sites and facilities appropriate with the production scale, product variety, and inspection requirements shall be equipped with. Verify whether the enterprise have appropriate inspection qualifications by taking reference to the production process requirements, the product inspection requirements and inspection methods. |
|  | 2.8.1 | A clean production environment shall be provided. The ground, road surroundings, and transportation in plant area shall not introduce contamination to the manufacturing. The plant shall stay away from air and water contamination and other contamination sources. Check on-site to ensure the production environment is clean, free of stagnant water and weed. The ground, road surroundings, and transportation in plant area shall not introduce contamination to the manufacturing. Check the flatness of the ground and road, measures for reduction of bare soil, flying dust, and greening, storage of garbage, unused items in the plant. |
|  | 2.8.2 | The general layout of administration, living and ancillary areas shall be well designed to avoid interference from each other, and shall not adversely affect the production area. |
|  | 2.9.1 | The production plant shall be equipped with dust-proof facilities and facilities to prevent entry of insects and other animals. Doors, windows and emergency doors of the clean room (area) shall be sealed, and the doors of the clean room (area) shall be opened towards the direction of higher cleanliness. |
|  | 2.10.1 | According to production process control for in vitro diagnostic reagents, determine the production process within clean rooms (areas) of appropriate levels to avoid contamination in the production. Check relevant documents, to see whether the cleanliness level of the production process has be defined; inspect on-site to see if production is carried out in a clean room (area) with the appropriate level, and whether contamination in the production process can be avoided. |
|  | 2.10.2 | Differential static pressure between differently classified clean rooms (areas) shall be greater than 5 Pa, differential static pressure between the clean room (area) and the outdoor air shall be greater than 10 Pa, and devices indicating the differential pressure shall be equipped with. Check on-site to see whether the differential pressure device has been equipped with, and whether the differential static pressure between clean rooms (areas) with different air cleanliness levels or between the clean room (area) and the outdoor air have met requirements. |
|  | 2.10.3 | Pressure gradient between clean rooms of the same level shall be reasonable. Check relevant documents to see whether the pressure gradient between clean rooms of the same level is clearly defined, check on-site to see whether the requirements are met. |
|  | 2.11.1 | For preparation and filling, such as dosing, coating, filling, spot foiling, drying, cutting, foiling and sub-packaging of ELISA reagents, immunofluorescence reagents, immune luminescent reagents, polymerase chain reaction (PCR) reagents, gold-labeled reagents, dry chemical reagents, cell culture media, calibrators and control materials, enzymes, antigens, antibodies and other active ingredients, the cleanliness level of the production areas shall not be lower than 100,000. |
|  | 2.12.1 | For processing operations of negative or positive serum, plasmid or blood products, the cleanliness level of the production areas shall not be lower than 10,000, and relatively negative pressure shall be maintained between the production area and the adjacent areas. |
|  | 2.13.1 | For filling of sterile materials, the cleanliness level of some parts of the operation area shall be 100. |
|  | 2.14.1 | Production of general chemical reagents shall be carried out in a clean environment. |
|  | 2.15.1 | Air cleanliness levels of the clean room (area) shall conform to the requirements in the following table: Table of air cleanliness levels of clean rooms (areas) Cleanliness level; Maximum number of dust particles/m3; Maximum number of microbes. ≥0.5μm; ≥5μm; Planktonic bacteria/m3; Settling microbes /dish. Level 100; 3,500; 0; 5; l. Level 10,000; 350,0O0; 2,000; 100; 3. Level 100,000; 3,500,000; 20,000; 500; 10. |
|  | 2.16.1 | The clean room (area) shall be rationally arranged in accordance with production process and the required cleanliness level of in vitro diagnostic reagents, and the directions of pedestrian flow and logistics shall be reasonable. Production operation cross-contamination in the same clean room (area) between adjacent clean rooms (areas) shall be avoided. Check on-site to see whether the directions of pedestrian flow and logistics are reasonable, and whether cross-contamination can be avoided. |
|  | 2.17.1 | Entrances to the pipes, air inlet and outlet of the clean room (area) shall be rationally arranged, interfaces between water, electricity and gas transmission lines and the wall shall be reliably sealed, and lighting shall not be suspended. |
|  | 2.18.1 | Temperature and relative humidity of the clean room (area) shall be commensurate with the production process. Without special requirements, the temperature shall be controlled at 18~28℃, and the relative humidity shall be controlled in 45%~65%. Check on-site to see whether the temperature and humidity devices and recording have met the requirements. |
|  | 2.19.1 | Buffer facilities shall be provided between the clean room (area) and the non-clean room (area). |
|  | 2.20.1 | Interior surfaces (walls, floors, ceilings and consoles) of a clean area shall be smooth, free from cracks, open joints and dust retention, shall not shed particulate matter, and shall permit easy and effective cleaning and disinfection.. |
|  | 2.21.1 | If air in the clean room (area) is reused, effective measures shall be taken to prevent contamination or cross-contamination. Check air supply, return and exhaust drawings of the air conditioning and cleaning system, if the air is reused, verify the effectiveness of the control measures. |
|  | 2.22.1 | Devices shall be installed with sinks, floor drains in the clean room (area) to prevent flow backward, so as to avoid contamination to the environment and materials. |
|  | 2.22.2 | Floor drain shall not be set in a clean room (area) with the cleanliness level of 100. |
|  | 2.23.1 | Operation areas where dust is generated should be kept under relatively negative pressure, and specific measures shall be taken to avoid dust diffusion and cross-contamination. |
|  | *2.24.1 | Polluting, infectious materials and highly active materials shall be processed under controlled conditions to avoid infection, contamination or leakage. |
|  | *2.25.1 | For production of hormones, operations of pathogenic pathogens or bacillus products, separate air purification systems shall be used, negative pressure shall be maintained with the adjacent areas, and the exhausted air shall not be reused. Check on-site to see whether the production areas of hormones, operating areas of pathogenic pathogens or bacillus products are strictly separated from production areas or other products. If it is a multistory building, in the same story, the production lines shall not share material channel, personnel access, packaging line with production lines of other general production lines, to avoid cross-contamination. Check on-site to see whether there is an independent air purification system, and ensure the air exhausted can not be reused. Check on-site the differential pressure gauge, to see whether negative pressure is maintained with the adjacent areas. |
|  | 2.26.1 | Biological safety cabinets shall be provided for operations of pathogens with danger degree of grade II or above, and the air shall be filtered before exhausting. |
|  | 2.26.2 | Performance of the filter shall be checked on a regular basis to ensure its effectiveness. Check the maintenance records, to see whether regular checks are carried out. |
|  | 2.26.3 | Appropriate protective measures shall be taken for positive serum using pathogen detection reagents. Check relevant documents to see whether protective measures are provided. Check on-site to see the storage conditions and use records of positive serum. |
|  | *2.27.1 | For collection and preparation of high pathogenic pathogens, appropriate bio-safety facilities shall be equipped with in accordance with the industry standards issued by relevant departments, such as Human Infectious Disease and Microorganisms Directory, Microbiological and Biomedical Laboratories General Biosafety Guidelines, laboratory bio-safety-General requirements and other relevant provisions. Check relevant documents, to see whether the danger degree of pathogens, the biosafety requirements are in conformity with the relevant provisions. Check on-site to see whether bio-safety facilities are equipped with, and whether the protection requirements are met. |
|  | *2.28.1 | For polymerase chain reaction (PCR) reagents, the production and detection shall be carried out in a separate building or space, and the air shall not directly circulated, to prevent cross contamination from aerosol formed during the amplification. For polymerase chain reaction (PCR) reagents, the production and inspection apparatus shall not be mixed, and shall be strictly cleaned and disinfected. |
|  | 2.29.1 | Number of personnel in the clean room (area) shall be commensurate with the floor space of the clean room (area). Check the validation records to see whether validation has been carried out on the upper limit of the number of personnel on-site, to ensure the cleanliness control requirements are met. Verify the number of personnel on-site, check relevant records, and the number of personnel on-site shall not exceed the upper limit confirmed. |
|  | 2.30.1 | In vitro diagnostic reagents without air purification requirements on the production environment, production shall be carried out in a clean environment. Basic requirements on clean conditions: shall have dust-proof, ventilation measures, and other measures to prevent entry of insects or other animals, as well as foreign matters; pedestrian flow and logistics shall be separated, cleaning measures, such as changing shoes, clothes, wearing masks and hats, hand washing, hand disinfection other cleaning measures shall be taken before access to the workshop; floor of the production site shall be easy to clean, walls and ceiling shall be flat and smooth, without shedding particles; consoles shall be smooth, flat, seamless, corrosion-resistant, easy to clean and disinfect; the production areas shall be regularly cleaned, washed and disinfected; temperature and humidity of the workshop shall be controlled according to the production requirements. |
|  | *2.31.1 | Management of flammable, explosive, toxic, harmful, polluting or infectious materials, materials of highly active or materials derived from organisms shall comply with the relevant national regulations. Materials involved shall be listed, stored in special areas, safe kept and released by special personnel, and appropriate protective procedures shall be developed. Check the list to see whether all flammable, explosive, toxic, harmful, polluting or infectious materials, materials of highly active or materials derived from organisms have been included. Check the protective procedures to see whether the management and protection requirements are clearly defined. Check on-site to see whether the materials are safe kept in special areas and by special personnel, and whether there are releasing records. |
|  | 2.32.1 | Animal room shall be located in a building with good isolation, and shall be separated from the production and quality inspection areas to avoid contamination. |
| Equipment | *3.1.1 | Production equipment and process equipment commensurate with the production of products and scale shall be equipped with, and their effective operations shall be ensured. Check the equipment list by taking reference to the production flow chart to see whether the equipment listed has met the production needs; verify to see whether the field equipment is consistent with related content of the equipment list; equipment management system shall be established. |
|  | 3.2.1 | Design, selection, installation, repair and maintenance of the production equipment shall be consistent with its intended use, and shall be easy to operate, clean and maintain. Check the equipment validation records to confirm whether the requirements are met. Check on-site to see whether the equipment is easy to operate, clean and maintain. |
|  | 3.2.2 | The production equipment shall be clearly identified to prevent unintended use. Check the equipment identification on site. |
|  | 3.2.3 | Operating instructions for the use, cleaning, maintenance, and repair of production equipment shall be established, and corresponding equipment operation records shall be kept. |
|  | *3.3.1 | Appropriate inspection instruments and equipment that meet the product detection requirements shall be equipped with, and clear operating instructions shall be provided for major inspection instruments and equipment. Verify to see whether the enterprise has appropriate inspection equipment by taking reference to the product inspection requirements and methods. Whether operating instructions have been provided for major inspection equipment. |
|  | 3.4.1 | Use records of inspection instruments and equipment shall be established, of which the contents shall include the use, calibration, maintenance, repair, and other status, etc. |
|  | 3.5.1 | Appropriate measuring instruments shall be equipped with, the range and accuracy of the measuring instruments shall meet the requirements, the calibration validity date of the measuring instruments shall be indicated, and corresponding records shall be kept. Check the measuring instrument calibration records to determine whether they are used within the validity period. |
|  | 3.6.1 | Air purification system of the clean room (area) shall be subject to written confirmation and shall be maintained in continuous operation and with appropriate level of cleanliness, which shall be re-verified after a certain period. Check the verification and re-verification records of air purification system of the clean room (area). |
|  | 3.6.2 | If the air purification system is restarted after being shut down, necessary tests or verification shall be conducted to confirm whether the required cleanliness level are met. If air purification system of the clean room (area) is not in continuous use, validation shall be made to clarify its re-enable requirements, and operating records of each use of the air purification system shall be checked to see whether the control requirements are met. If validation is not carried out, environment parameters of the clean room (area) shall be detected before starting production again to confirm relevant standards are met. |
|  | 3.7.1 | The process water required shall be identified, when process water is used in the production process, appropriate equipment for preparing water shall be equipped with, and measures to prevent contamination shall be taken. If the water consumption amount is large, it shall be piped to the water using point in the clean room (area). Process water shall meet the requirements of product quality. Check on-site to see whether the water preparing equipment has met the water requirements; and whether measures to prevent contamination have been taken. |
|  | 3.8.1 | Process water management documents shall be established. Check the process water management documents, to see whether the process water types are in compliance with standard requirements of the "Pharmacopoeia" or GB/T6682 or YY/T1244, etc. Whether the equipment maintenance, repair, cleaning, disinfection, water quality monitoring and detection requirements are included. |
|  | 3.8.2 | Process water tanks and pipelines shall meet the requirements of the products on water quality, and shall be cleaned and disinfected on a regular basis. Check on-site to confirm that the process water tanks and pipelines are made of stainless steel or other nontoxic materials, are cleaned and disinfected on a regular basis, and corresponding records are made. |
|  | 3.9.1 | Material name, flow direction of the main fixed pipeline that connecting the ingredients tank and the equipment shall be indicated, the pipeline shall be cleaned and maintained on a regular basis, and the operation status of the equipment shall be indicated. Check on-site to see whether the label of the main fixed pipeline is clear, the material name and flow direction are indicated, and whether the operation status of the equipment connected to it is indicated. Check the cleaning and maintenance records of the main fixed pipeline. |
|  | 3.10.1 | Surfaces of equipment, containers and pipelines in direct contact with the materials or products shall be smooth, flat, free of shedding particles, nonpoisonous, corrosion-resistant, shall not react with or adhere to materials or products, and shall be easy to clean, treat, disinfect or sterilize. |
|  | 3.11.1 | Raw materials, semi-finished products, and finished products need to be refrigerated shall be equipped with appropriate cold storage equipment, of which the operation status shall be monitored, and the storage temperature shall be recorded. Check relevant documents to see whether the storage conditions of raw materials, semi-finished products and finished products are specified. Check on-site to see whether cold-storage equipment and facilities, such as refrigeration house, are equipped with, and whether the refrigeration house is provided with temperature indicator, alarm function, and emergency measures against power failure. Check the temperature records of the storage environment, to see whether the product storage requirements are met. |
|  | 3.11.2 | Refrigerated in vitro diagnostic reagents shall be equipped with transport facilities and equipment that meet the temperature requirements; Check the transportation management documents of refrigerated in vitro diagnostic reagents to see whether the protection requirements and methods of the reagents during transport are specified. Check on-site to see whether appropriate transport facilities and equipment are equipped with, and whether the temperature requirements for storage and transportation are met, |
| Documentation Management | *4.1.1 | Sound quality management system documents, including quality policy and quality objectives, quality manual, procedures, technical documents and records, and other documents required by regulations shall be established. The quality policy shall be communicated and understood within the enterprise; and shall be reviewed in terms of continuing suitability. Quality objectives shall be consistent with the quality policy; shall be decomposed into the quality objectives of all functions and levels based on the total quality objective; shall include the content required to meet product requirements; shall be measurable and assessable; shall be ensured by specific methods and procedures. |
|  | 4.1.2 | The quality manual shall made provisions for the quality management system. Check the quality manual of the enterprise, to confirm that it has included the quality objectives, organizations and responsibilities of the enterprise, scope of application of the quality system and requirements. |
|  | 4.1.3 | The procedures shall be developed based on various operating procedures needed to be established in the production and quality management process, and shall include various procedures contained herein. |
|  | *4.1.4 | Technical documents shall include the technical requirements and related standards, production instructions, operating instructions, inspection and testing procedures, installation and service procedures, and other related documents. |
|  | 4.2.1 | Document control procedures shall be established, to systematically design, prepare, review, approve and distribute the documents. |
|  | 4.2.2 | Operation procedures shall be followed for drafting, revision, review, approval, replacement or withdrawal, reproduction, storage and destruction of documents; and relevant records for distribution, withdrawal, reproduction and destruction of documents shall be in place.. |
|  | 4.2.3 | The update or revision of documents shall be reviewed and approved in accordance with provisions, and the change and revision status of the documents shall be able to be identified. Check relevant records to see whether update and revision of documents are reviewed and approved; and whether their change or revision status is identified. |
|  | 4.2.4 | Text of the document shall be suitable for distribution and use, documents have been revoked or canceled shall be identified to prevent misuse. Check the documents used in the working field to confirm whether they are valid versions. Whether obsolete documents are clearly identified. |
|  | 4.3.1 | Retention period of obsolete technical documents and other necessary quality management system documents shall be specified to meet the needs of product maintenance and traceability of product quality responsibility. The retention period shall not be less than the medical device life cycle of the enterprise . |
|  | 4.4.1 | Record control procedures shall be established, including records of identification, storage, retrieval, retention time and disposition requirements, etc. |
|  | 4.4.2 | Records shall ensure the traceability of activities such as production and quality control. |
|  | 4.4.3 | Records shall be clear and complete, readily identifiable and retrievable, and shall be protected against damage and loss. |
|  | 4.4.4 | Records shall not be altered or destroyed, names and dates shall be endorsed to changes in records, and the original message shall still be clearly visible, if necessary, the reason for this change shall be indicated. |
|  | 4.4.5 | Records retention period shall at least be equivalent to the life cycle of medical devices produced by the enterprise, and shall not be less than 2 years from the release date of the product, or in accordance with the relevant regulatory requirements, and shall be traceable. |
| Design and development | 5.1.1 | Design control procedures shall be established and documented, and planning and control for the design and development process of medical devices shall be implemented. Check to confirm that the design control procedures are clear, actionable, is able to control the design and development process, and at least includes the following contents: 1. Division of design and development at all stages; 2.Review, verification, confirmation and design transformation activities appropriate to each design and development stage; 3. Personnel and departmental responsibilities, rights and communications at all stages of design and development; 4. Risk management requirements. |
|  | 5.2.1 | When carrying out design and development planning, all the design and development stages and the review, validation, confirmation and design transformation activities of each stage shall be determined, the design and development activities interfaces of various departments shall be identified and determined, and the functions and division of work shall be clarified. Check the design and development planning information. Design and development activities shall be planned in accordance with the product features, and the planning results shall be documented, which shall include the following contents: 1. Description of the objectives and significance of the project being designed and developed, and analysis of technical indicators; 2. Determination of all stages of the design and development, and appropriate review, validation, confirmation and design transformation activities for each stage; 3.The design and development activities interfaces of various departments shall be identified and determined, personnel or organizational responsibilities and composition of reviewers, as well as the expected output of each stage shall be clarified; 4. The planning and arrangement of the main task and short-term tasks are consistent with the entire project; 5. Determination of the required measuring devices for the development, verification, confirmation of technical requirements and production activities; 6. Risk management activities. Design and development shall be carried out in accordance with the plan, of which when modification is required due to deviation of the design and development, the plan shall be reviewed and re-approved. |
|  | 5.3.1 | Design and development inputs shall include functionality, performance, and safety requirements provided by the intended use, regulatory requirements, risk management controls and other requirements. |
|  | 5.3.2 | Design and development inputs shall be reviewed and approved, of which related records shall be maintained. |
|  | *5.4.1 | Design and development outputs shall meet the input requirements, including procurement, production and service-related information, product specifications and so on. Check the information on design and development outputs, to confirm that it has complied with the following requirements: 1. Procurement information, such as technical requirements of raw materials, packaging materials, components and parts; 2. Information required for production and services, such as product drawings (including part drawings), technical formulas, work instructions, and environmental requirements, etc.; 3. Product specifications; 4. Product inspection procedures or instructions; 5. Provision of product features required for the safe and proper use of the product, such as product brochures, packaging and labeling requirements, etc. Whether the product manual is consistent with the registration and approval; 6. Identification and traceability requirements; 7. Files submitted to the registration and approval departments, such as research data, product technical requirements, registration inspection report, clinical evaluation data (if any), basic requirements on safety and effectiveness of medical devices, etc.; 8. Model machines or samples; 9. Biological evaluation results and records, including the main performance requirements of materials.. |
|  | 5.4.2 | Design and development outputs shall be approved, of which related records shall be kept. |
|  | 5.5.1 | Transformation activities from design and development to production shall be carried out in the design and development process to enable the validation of design and development outputs before forming final product specifications, so as to ensure that the design and development outputs are suitable for production. Check relevant documents to confirm that at least the following requirements are met: 1. Design transformation activities shall be carried out in the design and development process to address the problems of manufacturability, availability of parts and materials, required equipment, training of operating personnel, etc; 2. Design transformation activities shall correctly transfer each technical requirement of the product into specific processes or procedures associated with the product realization; 3. Records of design transformation activities shall show that the design and development outputs have been verified before forming final product specifications, and the validation records shall be kept, to ensure that the design and development outputs are suitable for production; 4. The transformation of special process shall be confirmed, to ensure that is results are applicable to the production, and the confirmation records shall be kept. |
|  | 5.6.1 | Review shall be scheduled at appropriate stages of the design and development, and records of assessment results and any necessary measures shall be kept. Check relevant documents and records, to ensure that the following requirements are at least met: 1. Design and development review shall be carried out at the appropriate stages according to the results of design and development planning; 2. Design and development review records shall be kept, including records of the review results and the necessary measures taken by the review. |
|  | 5.7.1 | Design and development shall be verified to ensure that the design and development outputs have met the input requirements, and records of the validation results and any necessary measures shall be kept. Check relevant documents and records, to ensure that the following requirements are at least met: 1. Design and development validation shall be carried out at the appropriate stages according to the results of design and development planning, to ensure that the design and development outputs have met the input requirements; 2. Records of the design and development validation, the validation results and any necessary measures shall be kept; 3. If alternative calculation methods or comparison of proven designs are used for design and development validation, the appropriateness of the methods used shall be reviewed to confirm that it is scientific and effective. |
|  | 5.8.1 | Design and development validation shall be carried out to ensure the products meet the requirements of specified application or intended use, and records of the validation results and any necessary measures shall be kept. Check relevant documents and records, to ensure that the following requirements are at least met: 1. Design and development validation shall be carried out at the appropriate stages, to ensure that the product meets the requirements of specified application or intended use; 2. Design and development validation activities shall be carried out before the delivery and implementation of the product; 3. Design and development validation records, including clinical assessment or clinical trial records, as well as records of the validation results and any necessary measures shall be maintained. |
|  | 5.9.1 | Confirm that clinical evaluation or performance evaluation is available. Clinical trials shall meet the regulatory requirements for clinical trials of medical devices . Check the clinical evaluation reports and supporting materials. If clinical trials are carried out, they shall comply with regulatory requirements and appropriate documentation shall be provided. For medical devices require clinical or performance evaluation, evaluation reports and (or) materials shall be available. |
|  | 5.10.1 | Design and development changes shall be identified and records thereof shall be maintained. |
|  | 5.10.2 | If necessary, changes to the design and development shall be reviewed, verified and confirmed, and shall be approved before implementation. Check review records of design and development changes, to ensure that the following requirements are at least met: 1. Impacts of changes on product components and products that have been delivered shall be included; 2. Design and development changes shall be implemented in accordance with relevant provisions of the registration of a medical device; 3. If the design changes and results involve changes in the contents of product registration certificates of medical devices (filing documents), the enterprise shall conduct a risk analysis and apply for change of registration (filing) in accordance with the provisions of relevant regulations, so as to meet regulatory requirements. |
|  | *5.10.3 | If changes in the selection of materials, parts or production functions may affect product safety and effectiveness of medical devices, the risks brought by such changes shall be evaluated, and if necessary, measures shall be taken to reduce the risk to an acceptable level, meanwhile, relevant regulatory requirements shall be in line with. |
|  | 5.11.1 | Risk management requirements shall be developed and documented in the process of product realization which includes design and development, and relevant records shall be maintained. Check the risk management documents and records, to ensure that the following requirements are at least met: 1. Risk management shall cover the entire process product realization development. 2. Risk management documents for medical devices shall be established, and related records shall be maintained to identify evidence of implementation. 3. Risks of medical devices shall be controlled to acceptable levels. |
|  | 5.12.1 | Development conditions, including use of equipment, instruments and reagents shall meet the development needs, and records the equipment, instruments and reagents used shall be kept. |
|  | 5.12.2 | In the development process, main sources of raw materials, intermediates, important excipients shall be clear. Records of their number, the amount used and the remaining amount shall be kept. |
|  | 5.12.3 | Records of sample quantities, storage conditions, samples, use or destruction status at all stages, such as technology study, technical requirements / analytical performance study, stability study, testing, clinical trials / evaluation (including pre-experimental) study, and reference value study shall be kept, and the quantity of trial-produced samples shall meet the requirement for study needs. |
| Procurement | *6.1.1 | Procurement control procedures shall be established. Procurement procedures shall at least include: procurement process, qualified supplier selection, assessment and re-assessment requirements, procurement of goods inspection or certification requirements, and the procurement records requirements. |
|  | *6.1.2 | It shall be ensured that the procurement of goods comply with specified requirements, and shall not be lower than relevant provisions of the laws and regulations, or relevant standards of mandatory requirements of the State. |
|  | 6.2.1 | The manner and extent of exerting control over procurement of goods shall be determined according to the impact of the procurement of goods on products. Check the provisions for manner and extent of exerting control over procurement of goods, to verify whether the control manner and control extent can meet product requirements. |
|  | 6.3.1 | Supplier audit system shall be established in order to audit and evaluate the suppliers. If necessary, on-site audit shall be conducted, to see whether it is in compliance with the Guidelines for medical device manufacturing enterprise supplier auditing requirements. |
|  | 6.3.2 | Records of supplier evaluation results and the evaluation process shall be kept. |
|  | *6.4.1 | Quality agreements shall be signed with the main suppliers of raw materials, to identify the quality responsibilities of both parties. |
|  | 6.5.1 | When purchasing, the purchasing information must be made clear, and the procurement requirements must be described in detailed, including categories of procurement of goods, acceptance criteria, specifications, procedures, drawings and so on. Check from the list of purchased items the procurement requirements, to confirm compliance with the requirements of this article. |
|  | 6.5.2 | Procurement records shall be established, including procurement contracts, inventories of raw materials, suppliers qualification documents, quality standards, inspection reports and acceptance standards, etc. |
|  | *6.5.3 | Procurement records must meet the traceability requirements. |
|  | 6.6.1 | Procurement of goods shall be inspected and verified to ensure that the production requirements are met. Check the inspection or verification records of procurement of goods. |
|  | *6.7.1 | Procurement of out-sourced standards, calibrators, controls, and blood products used for control shall satisfy the traceability requirements. |
|  | *6.7.2 | The enterprise or the supply agency shall be responsible for determination of the pathogenic microorganisms and explicitly setting of the value range; clear records of their sources, value ranges, inactive state, quantities, storage and use state and other information shall be kept, and shall be taken care by special personnel. Check the appropriate documents, to see whether the record requirements and the requirements of the person responsible are specified. Check related records of the pathogenic microorganisms, to see whether the requirements are met. |
| Production Management | *7.1.1 | Production shall be carried out in accordance with the quality management system established, to ensure compliance with mandatory standards and the registered or filed product technical requirements. |
|  | *7.2.1 | Production process and operating instructions shall be prepared, to identify key processes and special procedures. Check related documents; whether key processes and special procedures are specified, whether validation or confirmation requirements on important parameters for key procedures and special procedures are provided. |
|  | 7.3.1 | In the production process, if the raw materials and intermediates need to be cleaned, the clean methods and requirements shall be specified, and the cleaning effect shall be verified. |
|  | 7.4.1 | The environment shall be monitored according to the production process characteristics, and related records shall be kept. |
|  | 7.5.1 | Special processes of the production shall be confirmed, and related records shall be kept, including the confirmation plan, confirmation method, operators, result evaluation, and the reconfirmation, etc. |
|  | 7.5.2 | If the computer software used in the production process have an impact on the product quality, verification or confirmation shall be carried out. |
|  | *7.6.1 | Products of each batch (set) shall have a production record and shall meet the requirements of traceability. |
|  | 7.6.2 | Production records shall include: product name, specifications, raw material batch number, production batch number or product number, production date, quantities, major equipment, process parameters, and operating personnel, etc. |
|  | 7.7.1 | Product identification procedures shall be established, using proper methods to identify the product, in order to identify and prevent mixed use and misuse. |
|  | *7.8.1 | The test status of the product shall be identified in the production process, to prevent rejected intermediates flowing to the next procedure. Check to see whether the test status identification methods are provided, check on-site the test status identification in the production process to see whether it is in compliance with the provisions of the document. |
|  | *7.9.1 | A procedure for product traceability shall be established, providing the traceability scope, extent, identification and required records. |
|  | *7.10.1 | Product manuals, labels shall be in conformity with the relevant laws, regulations and standards. |
|  | 7.11.1 | Product protection procedures shall be established, providing the protection requirements of the product and its components, including contamination prevention, ESD protection, dust protection, corrosion protection, transport protection, and other requirements. Protection shall include identification, handling, packaging, storage and protection. Check on-site to see whether the product protection procedures are in line with standard requirements; check on-site and inspect randomly related records to confirm the product protection has met requirements. |
|  | 7.12.1 | The compressed air used in the clean room (area) and other process gases shall be purified. Check on-site the purification devices and piping of the process gases, and confirm whether the process gases have been purified. |
|  | 7.12.2 | For gases in direct contact with the use surfaces of the product, their effects on the product shall be verified and controlled, in order to meet the product requirements. Check the evaluation and validation records of effects of gases in direct contact with the use surfaces of the product on the product, to see whether control measures are provided and implemented in accordance with the evaluation and validation results. |
|  | 7.13.1 | Production equipment, containers shall conform to the clean environment control and process document requirements. Check relevant documents of the production equipment and containers, to see whether their cleaning requirements have been provided, and whether they have met the requirements of environmental control and process documents |
|  | 7.14.1 | Classification and storage management shall be carried out in accordance with the properties and storage requirements of the materials, and the storage conditions and time limits of the intermediates shall be clearly defined. Check relevant documents, to see whether the materials are classified, and whether the requirements of classified storage, storage conditions and time limit of the intermediates are clarified. Check on-site to see whether the materials are stored by category, and whether their storage requirements are met. Check on-site to see whether the storage conditions and time limits of intermediates are in compliance. |
|  | 7.14.2 | Materials shall be used within the prescribed period according to the first in first out principle. For materials without useful life, the storage period shall be determined based on their stability data. Check the material management-related documents, to see whether the use principle of first in first out is clarified, whether storage period of materials without useful life is determined based on the stability data. Check the material requisition records to see whether the first in first out principle is complied. For materials without useful life, check the stability data. |
|  | 7.14.3 | In case changes in the storage conditions that might affect the product quality are found during the storage period, reinspection shall be carried out in a timely manner. Check the material storage environment records, and if the storage requirements are not met, check the reinspection records. |
|  | 7.15.1 | Materials entering the clean room (area) shall be purified in accordance with the procedure. Check relevant documents, to see whether relevant procedures for materials entering the clean room (area) are specified. Check on-site to see whether the purification procedure of the materials entering the clean room (area) is in conformity with provisions. |
|  | 7.16.1 | In the production process, product identification and production status identification control procedure shall be established. |
|  | 7.16.2 | Identify and manage status of all kinds materials, production areas, equipment, and pipelines on site. |
|  | 7.17.1 | Material balance checks shall be carried out on key raw materials of each batch of products. If there are significant differences, the root causes must be identified to come up with a reasonable explanation, and after making sure that there are no potential quality accidents, the materials can be processed as normal products. |
|  | 7.18.1 | Batch management system shall be established. Batch management shall be carried out on key raw materials, intermediates and finished products according to the provisions, and traceability records shall be maintained and provided. Check the batch management system, to see whether batch number coding methods of key raw materials, intermediates and finished goods are specified. Check relevant records, to see whether the key raw materials, and intermediates are traceable. |
|  | 7.18.2 | For components in the same kit with different batch numbers, try to combine components with the closest dates of production together, and the batch number and expiry date of each component shall be respectively indicated on its container. |
|  | 7.18.3 | The useful life of the entire kit shall be subject to the component first comes to its expiry date. |
|  | 7.19.1 | The production of different varieties of products shall be effectively isolated, in order to avoid mixing up and contamination. |
|  | 7.19.2 | When a number of packaging lines are operating at the same time, isolation or other effective measures shall be taken to prevent mixing up. |
|  | 7.20.1 | Clean room (area) hygiene management documents shall be developed. The clean room (area) shall be cleaned and disinfected in accordance with provisions , and records shall be kept. Check the clean room (area) process hygiene management documents and records, to see whether the following contents are contained: 1. equipment cleaning requirements; 2. station equipment (containers, pipes, tanks, etc.) cleaning requirements; 3. material cleaning requirements; 4. console, space, wall, ceiling cleaning requirements; 5. cleaning and storage provisions of cleaning tools ; 6. clean room (area) air disinfection requirements; 7. disinfectant selection, use regulations. Check on-site to see whether cleaning tools in clean room (area) have no sheddings, are easy to clean and disinfect, classified by use, and cleaning tools of different clean rooms (areas) are not be cross-used. |
|  | 7.20.2 | Disinfectants or disinfection methods used shall not cause contamination to equipment, containers, materials and products. Check relevant documents, to see whether disinfectants or disinfection methods are provided, and whether the provisions include the disinfectant types, usage time, frequency, replacement cycle and other contents, and relevant records shall be retained. Evaluation or validation shall be carried out on the disinfection methods or disinfectants selected; and the disinfectants or disinfection methods used shall not cause contamination to devices, containers, materials and products. |
|  | 7.20.3 | Disinfectant types shall be altered on a regular basis, to prevent the development of drug-resistant strains. Check on-site to see whether the disinfectants used are consistent with the document provisions, and whether they are altered on a regular basis. |
|  | 7.21.1 | Lubricants, cleaning agents used in the production equipment shall not pollute the product. Check the the validation reports, to ensure that the use of lubricants, cleaning agents will not pollute the product, or appropriate measures are taken to eliminate the effects of contamination. |
|  | 7.22.1 | Site-clearing regulations shall be established. Whether management provisions and records on site-clearing are made according to the production process. |
|  | 7.22.2 | Site-clearing shall be carried out after the previous process is finished or production of the previous product is completed, other production shall not be started before qualified acceptance, and related records of sit-clearing shall be kept. Check on-site to see whether disposal of remnants of the last production process comply with the requirements, and check the site-clearing records. |
|  | 7.22.3 | Preparation and filling appliances must be dedicated, and shall be cleaned, dried and otherwise handled after use. |
|  | 7.23.1 | Traceability procedures shall be established and documented, in which the scope and degree of traceability, identification and recording shall be provided. |
|  | 7.23.2 | Records shall include the raw materials used in the production process, the production processes, production equipment, operators and production environments, etc. |
|  | 7.24.1 | After a certain period, key projects shall be reverified. Check the validation documents, to see whether revalidation requirements of key projects are specified, and check the related records. |
|  | 7.24.2 | When conditions for revalidation of the main factors affecting the quality of products, such as technology, quality control methods, main starting materials, main production equipment are changed, revalidation shall be carried out on related content. Check related validation documents, to see whether revalidation requirements for changes in these factors are provided, and check the related records. |
|  | 7.24.3 | The enterprise shall propose the validation time based on different product characteristics. |
|  | 7.25.1 | If production has been discontinued for more than 12 months, before reorganizing production, revalidation on the production environment and facilities, equipment, main raw materials, key processes, test equipment and quality control methods shall be carried out. Check related documents and records, to see whether the requirements aremet. |
|  | 7.25.2 | In case of stop of production is less than 12 months, if necessary, the enterprise shall also reverify the production environment and the facilities and equipment. |
|  | 7.26.1 | Inactivation procedures shall be established for serum or plasma used for production which need to be inactivated. Check the inactivation procedures, to see whether the types of serum or plasma used for production which need to be inactivated, the inactivation methods, protective measures in operation, isolation and identification requirements of the materials before and after inactivation are specified. Check the inactivation records. |
|  | 7.26.2 | According to the requirements of the procedures, the state of serum or plasma before and after the inactivation used for production shall be clearly distinguished and identified. Check on-site to see whether the serum or plasma before and after the inactivation are distinguished and identified, and the identification shall include the name of the serum or plasma, batch number, inactivation state and inactivation date. |
|  | 7.27.1 | Harmless treatment shall be carried out for waste liquids and waste substances produced in the production process, and such treatment shall be in conformity with the relevant requirements of environmental protection. |
| Quality Control | 8.1.1 | Quality control procedures shall be established, providing product inspection department, personnel, operations, and other requirements. Check the quality control procedures, to see whether responsibilities of inspection departments, personnel qualifications, inspection procedures, etc. are provided. |
|  | 8.1.2 | Use and calibration requirements of inspection instruments and equipment, and product release procedures shall be provided. Check the quality control procedures, to see whether the use and calibration requirements of inspection instruments and equipment are provided. |
|  | 8.2.1 | Inspection instruments and equipment shall periodically calibrated or verified, and identified. Check the inspection instruments and equipment to see whether calibration and inspection are are carried out in accordance with provisions, and whether they are identified. |
|  | 8.2.2 | Protection requirements for inspection instruments and equipment during handling, maintenance, and storage shall be provided, to prevent inaccurate inspection results. |
|  | 8.2.3 | If the inspection instruments and equipment are found not meeting requirements, the inspection results in the past shall be evaluated, and the verification records shall be kept. Check the equipment use and maintenance records, to see whether the previous inspection results are evaluated in case of inspection instruments and equipment inconformity with requirements, and keep relevant records. |
|  | 8.2.4 | Computer software used for inspection shall be verified. |
|  | *8.3.1 | Product inspection procedures shall be developed according to the mandatory standards and registered or filed product technical requirements, and appropriate inspection reports or certificates shall be issued. Check to see whether the product inspection procedures have included performance indicators of mandatory standards and registered or filed product technical requirements; and verify whether the inspection records can substantiate the product in conformity with requirements; check to see whether appropriate inspection reports or certificates are issued in accordance with inspection procedures and inspection results. |
|  | 8.3.2 | In principle, inspections of incoming inspection, process inspection and finished product inspection projects which required routine controls shall not be commissioned. For projects with higher requirements on the inspection conditions and equipment, if commissioned inspection are required, a qualified inspection agency shall be commissioned to carry out the inspection, in order to prove that the products comply with the mandatory standards and the registered or filed product technical requirements. |
|  | *8.4.1 | Products of each batch (set) shall have batch inspection records and shall meet the requirements of traceability. |
|  | 8.4.2 | Inspection records shall include inspection records of incoming inspection, process inspection and finished product inspection, inspection reports or certificates, etc. |
|  | *8.5.1 | Product release procedures, conditions and approval requirements shall be provided. Check the product release procedure, to see whether the release conditions and release approval requirements are specified. Authorized release personnel and their duties and powers shall be provided, and approval records shall be kept. |
|  | 8.5.2 | Release of the products shall be accompanied by quality certificates. |
|  | 8.6.1 | Sample retention management provisions shall be developed based on products and process characteristics, samples shall be retained according to the provisions, and sample retention and observation records shall be kept. |
|  | 8.7.1 | Value traceability procedures of calibrators and references shall be established. Assign values to calibrators and references of each batch production. Check the value traceability procedures of calibrators and references (see ISO17511) and value assignment records of calibrators and references. |
|  | *8.8.1 | Sources of bacterial and viral strains used for production and inspection shall be indicated, the acceptance, storage, keeping, use and destruction of bacterial and viral strains shall comply with the national regulations on storage and keeping of medical microorganism strains and the pathogenic microorganism laboratory bio-safety regulations. Systems of original seed lots, master seed lots and work seed lots shall be established for bacterial and viral strains used for production and inspection . |
|  | *8.9.1 | Original cell bank, master cell bank and working cell bank shall be established for cells used for production. Cell library archives and cell operation log shall be established. |
|  | *8.9.2 | For self-prepared antigens or antibodies, detailed records of the source and nature of raw materials used shall be kept, and shall be traceable. |
|  | 8.10.1 | Ledger and use records shall be established for standards, calibrators, and controls used in the inspection process. Their origins, batch numbers, useful life, traceability ways, main technical indicators, preservation status shall be recorded. |
|  | 8.10.2 | Carry out re-inspection for standards, calibrators and controls in accordance with provisions, and keep records. |
|  | *8.11.1 | The retained samples shall be stored under specified conditions. Sample ledger shall be established, to timely record the sample inspection information, and the sample inspection report shall indicate the sample batch number, useful life, inspection date, inspector and inspection results, etc. Check to see whether sample retention management system is provided, whether the sample quantities, specifications, storage conditions and inspection requirements are clarified. Check on-site to see whether the environment of the sample retention room (area) is in conformity with the requirements of the retained samples, whether the required environment monitoring equipment is equipped with, and whether records are made. Check the sample retention ledger and sample inspection report to see whether requirements are met. |
|  | 8.11.2 | After expiration of the retention, the sample inspection reports shall be summarized, analyzed and archived. |
| Sales and After-sales Service | *9.1.1 | Product sales records shall be established and meet the traceability requirements. |
|  | 9.1.2 | Sales records shall at least contain: medical device names, specifications, types, quantity, batch number, expiry date, date of sale, name, address and contact information of the purchase unit, and so on. |
|  | 9.2.1 | Direct sales of self-produced products or selection of medical device enterprises shall comply with the medical devices regulations and specifications. |
|  | 9.2.2 | If the medical device enterprise is found with existence of illegal business activities, immediate report shall be made to the local food and drug supervision and management departments |
|  | 9.3.1 | The enterprise shall have the after-sale service ability commensurate with the products, and shall establish sound after-sales service system. |
|  | 9.3.2 | After-sales service requirements shall be provided, and after-sale service records shall be established and shall meet the traceability requirements. |
|  | 9.4.1 | For medical devices need to be installed by the enterprise, the installation requirements and installation verification acceptance criteria shall be determined, and installation and validation records shall be established. |
|  | 9.4.2 | If the devices are installed and maintained by the use unit or other enterprises, the enterprise shall provide installation requirements, standards and maintenance parts, information, passwords, and guidances, etc. |
|  | 9.5.1 | The enterprise shall establish customer feedback processing procedures, to track and analyze customer feedback information. Check the procedure documents to see whether provisions have been made on the implementation of the above activities, and whether the customer feedback information is tracked and analyzed. |
| Control of Nonconforming Products | *10.1.1 | Nonconforming product control procedures shall be established to provide responsibilities and rights of nonconforming product control department and personnel. |
|  | *10.2.1 | Nonconforming products shall be identified, recorded, isolated, and reviewed, and appropriate processing measures shall be taken based on the review results. Check on-site to see whether the identification and isolation of nonconforming products are in compliance with the provisions of the procedure documents; check the nonconforming product processing records, to see if they are reviewed in accordance with the provisions of the documents. |
|  | 10.4.1 | If nonconformities are found after the products are sold, appropriate measures, such as recall and destruction measures shall be taken in a timely manner. Check on-site the processing measures for nonconforming products found after sold, to see whether recall and destruction measures are taken. |
|  | 10.4.1 | For rework of nonconforming products, the enterprise shall prepare rework control documents. Rework control documents shall include the contents of operating instructions, re-inspection and revalidation. Check the rework control documents, to see whether provisions on reworkable nonconforming products are made; check the rework activity records, to confirm that the requirements of rework control documents are met. |
|  | 10.4.2 | For non-reworkable nonconforming products, the enterprise shall establish relevant disposal system. |
| Adverse event monitoring, Analysis and Improvement | 11.1.1 | Relevant departments shall be designated responsible for receiving, investigation, evaluation and dealing with customer complaints, and related records shall be kept. Check documents relating to rights and responsibilities, to determine whether provisions of the above-mentioned activities are made. |
|  | *11.2.1 | The adverse medical device event monitoring system shall be established in accordance with the relevant regulatory requirements, to carry out adverse event monitoring and reevaluation, and keep related records. Check the adverse event monitoring system of the enterprise, to see whether the responsibilities of management personnel of suspected adverse events, reporting principles, reporting procedures and reporting time frame are provided, whether procedures and documents for implementing medical device reevaluation are developed, and have complied with regulatory requirements. Check relevant records, to confirm whether there is any adverse event, and whether implementation has been carried out as required. |
|  | 11.3.1 | Data analysis procedures shall be established to collect and analyze data related to the quality of products, adverse events, customer feedback and quality management system operations, validate product safety and effectiveness, and keep relevant records. Check the implementation records of data analysis, to see whether it is carried out in accordance with provisions of the procedure, and whether the statistical techniques are applied and the data analysis results are recorded. |
|  | 11.4.1 | Corrective action procedure shall be established to determine the cause of the problem, and the effective measures to be taken to prevent the problem from happening again. |
|  | 11.4.2 | Preventive action procedure shall be established, to determine the cause of potential problems and the effective measures to be taken to prevent such problems from happening again |
|  | *11.5.1 | For medical devices with potential risks, recall and other measures shall be taken in accordance with relevant regulatory requirements, and reports shall be made to related departments according to the provisions. |
|  | 11.6.1 | Product information informing procedures shall be established to timely inform the use unit, related enterprises and customers about changes in products, use and other complement information. |
|  | 11.7.1 | Quality management system internal audit procedures shall be established, providing the audit criteria, scope, frequency, participants, methods, recording requirements, evaluation of the effectiveness of corrective and preventive measures and other contents, to ensure that the quality management system meets the requirements of this specification. Check the internal audit procedures to see whether the above contents have been included. Check the internal audit information to see whether the internal auditing personnel are trained, the internal audit records are in conformity with requirements, whether corrective measures have been taken for the problems found in the internal audit, and whether the measures taken are effective. |
|  | *11.8.1 | Management reviews shall be carried out on a regular basis, for evaluating and auditing the quality management system, to ensure its continuing suitability, adequacy and effectiveness. Check the management review documents and records, to see whether the management review plans, management review reports and related measures for improvements are included, and whether the management review reports have include evaluation of regulatory compliance. Whether management reviews are carried out within a specified time, whether improvement measures are proposed, whether specific responsibilities and requirements are implemented, and whether the implementation is carried out as planned. |

Note:

1. Numbering method of the guiding principle is X1.X2.X3, wherein, X1 is the chapter order number, for example, the first digit X1 of 1.1.1 represents the chapter "Organizations and Personnel ", the first digit X1 of 2.1.1 represents the chapter "Premises and Facilities"; X2 represents the order number of clauses in the same chapter, for example, the second digit X2 of 1.1.1 represents the first clause of the chapter "Organizations and Personnel ", the second digit X2 of 1.2.1 represents the second clause of the chapter "Organizations and Personnel "; X3 is the order number of the detailed inspection guidance in the same clause, for example, the third digit X3 of 1.1.1 represents the first detailed inspection point of the first requirement of chapter "Organization and Personnel", the third digit X3 of 1.1.2 represents the second detailed inspection point of the first requirement of chapter "Organization and Personnel".

Numbering rules of other chapters are the same.

2. Inspection guiding principles of each chapter are divided into two parts, the first part is the inspection guidance developed in accordance with the provisions of "Good Manufacturing Practice for Medical Devices", and the second part is the inspection guidance developed in accordance with the provisions of "Appendix to Good Manufacturing Practice for Medical Devices: In Vitro Diagnostic Reagents".

Office of the China Food and Drug Administration, issued on Sept. 25, 2015