| ICS 19.080;71.040.10 CCS N 20;N 61 |  |
| --- | --- |

National Standard of the People’s Republic of China

GB 4793—2024

Partially replaces GB 4793.1-2007

Safety technical specification for electrical equipment for measurement, control and laboratory use

Only used for full-text disclosure system of national standards. This text is only for personal study and research. Without authorization, reproduction, distribution, compilation, translation or network transmission is prohibited, and infringement will be prosecuted. National Public Service Platform for Standards Information: https://std.samr.gov.cn

Issued on October 28, 2024	Implemented on November 1, 2026

| Issued by | State Administration for Market Regulation Standardization Administration of China |
| --- | --- |

Table of Contents

FOREWORD	V

Introduction	VI

1	Scope	1

2	Normative references	1

3	Terms and definitions	1

4	Environmental Conditions	1

4.1	Normal environmental conditions	1

4.2	Extended environmental conditions	2

5	Marking and documentation	2

5.1	Marking	2

5.2	Warning markings	5

5.3	Durability of markings	5

5.4	Files	5

6	Protection against electric shock	7

6.1	General rule	7

6.2	Exceptions	8

6.3	Limit values for ACCESSIBLE parts	9

6.4	Primary means of protection	11

6.5	Additional means of protection in case of SINGLE FAULT CONDITION	11

6.6	Insulation requirements	14

7	Protection against mechanical HAZARDS	23

7.1	General rule	23

7.2	Sharp edges	23

7.3	Moving parts	23

7.4	Stability	26

7.5	Provisions for lifting and carrying	26

7.6	Wall mounting	27

7.7	Expelled parts	27

8	RESISTANCE TO MECHANICAL STRESSES	27

9	Protection against the spread of fire	27

9.1	General rule	27

9.2	Eliminating or reducing the sources of ignition within the equipment	28

9.3	Containment of the fire within the equipment, should it occur	29

9.4	Limited-energy circuit	30

9.5	Requirements for equipment containing or using flammable liquids	31

9.6	Overcurrent protection	32

10	Equipment temperature limits and resistance to heat	32

10.1	Surface temperature limits for protection against burns	32

10.2	Temperatures of windings	33

10.3	Resistance to heat	33

11	Protection against HAZARDS from fluids and solid foreign objects	33

11.1	General rule	33

11.2	Cleaning	34

11.3	Spillage	34

11.4	Overflow	34

11.5	Battery electrolyte	34

11.6	Equipment RATED with a degree of ingress protection (IP code)	34

11.7	Fluid pressure and leakage	34

12	Protection against radiation (including laser source), sound pressure and ultrasonic pressure	35

12.1	General rule	35

12.2	Equipment producing ionizing radiation	35

12.3	Optical Radiations	36

12.4	Microwave radiation	36

12.5	Sonic and ultrasonic pressure	36

12.6	Laser sources	37

13	Protection against released gases and substances, explosions and implosions	37

13.1	Poisonous and injurious gases and substances	37

13.2	Explosion and implosion	37

14	Components and subassemblies	38

14.1	General rule	38

14.2	Motors	38

14.3	Overtemperature protection devices	38

14.4	Fuse holders	38

14.5	MAINS voltage selecting devices	38

14.6	Printed wiring boards	39

14.7	Circuits used to limit TRANSIENT OVERVOLTAGES	39

15	Protection by interlocks	39

15.1	General rule	39

15.2	Prevention of reactivating	39

15.3	Reliability	39

16	Test Method	39

Appendix A  (normative)  Insulation requirements not covered by 6.6	41

A.1	Insulation for mains circuits	41

A.2	Insulation in secondary circuits	46

A.3	Insulation in circuits not addressed in 6.6, Clause A.1 or Clause A.2	52

REFERENCES	57

Figure 1 - Acceptable arrangement of protective means against electric shock	8

Figure 2 – Capacitance level versus voltage in NORMAL CONDITION and SINGLE FAULT CONDITION	10

Figure 3 - Examples of binding screw assemblies	13

Figure 4 - Distance between conductors on an interface between two layers	17

Figure 5 - Distance between adjacent conductors along an interface of two inner layers	17

Figure 6 - Distance between adjacent conductors located between the same two layers	18

Figure 7 - Flow chart to explain the requirements for protection against the spread of fire	27

Figure 8- Baffle	29

Figure 9 - Area of the bottom of an ENCLOSURE to be constructed as specified in 9.3.2 c) 1)	29

Figure A.1 - Distance between conductors on an interface between two layers	42

Figure A.2 - Distance between adjacent conductors along an interface of an inner layer	43

Figure A.3 - Distance between adjacent conductors located between the same two layers	45

Figure A.4 - Example of recurring peak voltage	53

Table 1 - Symbols	4

Table 2 - Multiplication factors for CLEARANCES of equipment RATED for operation at altitudes up to 5 000 m		15

Table 3 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOLTAGE CATEGORY II up to 300 V	16

Table 4 - Test voltages for solid insulation in mains circuits of OVERVOLTAGE CATEGORY II Up to 300 V		18

Table 5 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOLTAGE CATEGORY II up to 300 V	19

Table 6 - CREEPAGE DISTANCES for secondary circuits	20

Table 7 - Minimum values for distance or thickness (see 6.6.3.4.2 to 6.6.3.4.4)	21

Table 8 - Protective measures against mechanical HAZARDS to body parts	23

Table 9 - Minimum maintained gaps to prevent crushing for different body parts	24

Table 10 - Maximum gaps to prevent access for different body parts	25

Table 11 – Acceptable perforation of the bottom of an ENCLOSURE	28

Table 12 – Limits of maximum available current	30

Table 13 - Values for overcurrent protection devices	30

Table 14 – Surface temperature limits in NORMAL CONDITION	32

Table 15 - Maximum temperatures for insulation material of windings	32

Table 16 Test methods	38

Table A.1 - Multiplication factors for CLEARANCES for equipment RATED for operation at altitudes up to 5 000 m	41

Table A.2 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOL TAGE CATEGORY II above 300 V	41

Table A.3 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOLTAGE CATEGORY Ⅲ	41

Table A.4 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOLTAGE CATEGORY IV	41

Table A.5 - Minimum values for distance or thickness of solid insulation	43

Table A.6 - Test voltages for solid insulation in mains circuits of OVERVOLTAGE CATEGORY II above 300 V		44

Table A.7 -Test voltages for solid insulation in MAINS CIRCUITS of OVERVOL TAGE CATEGORY Ⅲ			44

Table A.8 -Test voltages for solid insulation in MAINS CIRCUITS of OVERVOL TAGE CATEGORY IV			44

Table A.9 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOLTAGE CATEGORY II above 300 V	46

Table A.10 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOL TAGE CATEGORY Ⅲ	47

Table A.11 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOLTAGE CATEGORY Ⅳ	48

Table A.12 - CREEPAGE DISTANCES for secondary circuits	49

Table A.13 - Minimum values for distance or thickness (see A.2.4.2 to A.2.4.4)	50

Table A.14 - CLEARANCE values for the calculation of A.3.2	52

Table A.15 - Clearances for basic insulation in circuits having recurring peak voltages or working voltages with frequencies above 30 kHz	53

Table A.16 Test voltages based on clearances	55

# FOREWORD

This Standard is drafted in accordance with GB/T 1.1—2020 Directives for standardization—Part 1: Rules for the structure and drafting of standardizing documents.

This Standard partially replaces GB 4793.1-2007 Safety requirements for electrical equipment for measurement, control and laboratory use-Part 1: General requirements. Compared with GB4793.1-2007, in addition to structural adjustments and editorial changes, the main technical changes of this Standard are as follows:

--Deleted test methods in each chapter and added Test Methods (see Chapter 16);

--Deleted the test content in the original Chapter 4 (see Chapter 4 of the 2007 edition);

--Added requirements for protection against electric shock of solid insulation (see Chapter 6);

--Requirements for mechanical hazard risk assessment were added to the Mechanical Hazard Prevention, and new technical requirements were added to the clearance limit between moving parts (see Chapter 7);

--Deleted the content of test and measurement equipment in the original Chapter 16 (see Chapter 16 of the 2007 edition);

--Added insulation requirements not covered in 6.7 (see Annex A);

--Deleted the contents of the original Annex A ~ Annex G (see Annex A ~ Annex G of the 2007 edition).

Attention is drawn to the possibility that some of the elements of this document may be the subject of patent rights. The issuing authority shall not be held responsible for identifying any or all such patent rights.

This document is proposed by and under the jurisdiction of the Ministry of Industry and Information Technology of the People's Republic of China.

The issuance of previous editions of this document:

--First published as GB 4793.1-1995 in 1995 and first revised in 2007;

--This standard is the second revision.

# Introduction

This document aims to ensure the safety of electrical equipment for measurement, control and laboratory use, so as to guard the bottom line of safety.

This document is the basic safety requirements for electrical equipment for measurement, control and laboratory use.

This document is used in conjunction with GB/T 42125 series of standards. GB/T 42125 series standards contain the technical requirements and corresponding test methods in this document.

Safety technical specification for electrical equipment for measurement, control and laboratory use

# 1	Scope

This document stipulates the safety requirements for marking and documentation, protection against electric shock, protection against mechanical hazards, resistance to mechanical stresses, protection against the spread of fire, equipment temperature limits and resistance to heat, protection against hazards from fluids and solid foreign objects, protection against radiation, including laser sources, and against sonic and ultrasonic pressure, protection against liberated gases and substances, explosion and implosion, components and general assembly, protection by interlocks, etc.

This document applies to the design, production, inspection and use of electrical equipment for measurement, control and laboratory use (including industrial automation equipment).

# 2	Normative references

The content of the following documents constitutes the essential articles of this document through normative references therein. For dated references, only the edition cited applies. For undated references, the latest edition of the referenced document (including any amendments) applies.

| GB/T 4025 | Basic and safety principles for man-machine interface, marking and identification - Coding principles for indicators and actuators |
| --- | --- |
| GB/T 4208 | Degree of protection provided by enclosures (IP code) |
| GB/T 5169.16 | Fire hazard testing for electric and electronic products - Part 16: Test flames - 50 W horizontal and vertical flame test methods |
| GB/T 7247.1 | Safety of laser products - Part 1: Equipment classification and requirements |
| GB/T 19661 | (All parts) Safety requirements for nuclear instrumentations and systems |
| GB/T 20138 | Degrees of protection provided by enclosures for electrical equipment against external mechanical impacts (IK code) |
| GB/T 42125.1-2024 | Safety requirements for electrical equipment for measurement, control and laboratory use-Part 1: General requirements |
| IEC 60027 | (All parts) Letter symbols to be used in electrical technology |

# 3	Terms and definitions

For the purposes of this document, the terms and definitions defined in GB/T 42125.1-2024 are applicable to this document.

# 4	Environmental Conditions

## 4.1	Normal environmental conditions

The equipment shall be designed to be safe for use at least under the following conditions:

a)	Indoor use;

b)	The altitude is not greater than 2,000 m;

c)	Temperature between 5 °C to 40 °C;

d)	Maximum relative humidity 80 % for temperatures up to 31 °C decreasing linearly to 50 % relative humidity at 40 °C;

e)	MAINS supply voltage fluctuations up to ±10 % of the nominal voltage;

f)	TRANSIENT OVERVOLTAGES up to the levels of OVERVOLTAGE CATEGORY Ⅱ;

g)	Temporary overvoltages occurring on the mains supply;

h)	Applicable pollution degree of the intended environment (pollution degree 2 in most cases).

## 4.2	Extended environmental conditions

This standard applies to equipment designed to be safe not only in the environmental conditions specified in 4.1, but also in any of the following conditions for which the equipment is RATED by the manufacturer:

a)	Outdoor use;

b)	Altitude above 2,000m;

c)	Ambient temperatures below 5 °C or above 40 °C;

d)	Relative humidity above the levels specified in 4.1;

e)	MAINS supply voltage fluctuations exceeding ±10 % of the nominal voltage;

f)	Wet locations;

g)	Transient overvoltages up to the levels of overvoltage category III or IV (according to Annex A).

# 5	Marking and documentation

## 5.1	Marking

### 5.1.1	General rule

Equipment shall bear markings as specified in 5.1.2 to 5.2. Except for marking of internal parts, these markings shall be visible from the exterior, or be visible after removing a cover or opening a door without the aid of a TOOL, if the cover or door is intended to be removed or opened by an OPERATOR. Markings applying to the equipment as a whole shall not be put on parts which can be removed by an OPERATOR without the use of a TOOL.

For rack- or panel-mounted equipment, markings are permitted to be on a surface that becomes visible after removal of the equipment from the rack or panel.

Letter symbols for quantities and units shall be in accordance with IEC 60027 (all parts). Graphic symbols shall be in accordance with Table 1, if applicable. Graphic symbols shall be explained in the documentation.

Note: Documentation include installation manual, operation manual, maintenance instructions, etc.

### 5.1.2	Identification

The equipment shall, as a minimum, be marked with:

a)	The name or trademark of the manufacturer or supplier.

b)	A model number, name or other means to identify the equipment. If equipment bearing the same distinctive designation (model number) is manufactured at more than one location, equipment from each manufacturing location shall be marked so that the location can be identified.

### 5.1.3	MAINS supply

The equipment shall be marked with the following information:

a)	Mains supply nature:

1)	a.c.rated mains frequency or range of frequencies;

2)	d.c.: symbol 1 of Table 1.

b)	Rated supply voltage(s) or range

c)	The maximum rated power in watts (active power) or volt-amperes (apparent power), or the maximum rated input current, with all accessories or plug-in modules connected. If the equipment can be used on more than one voltage range, separate values shall be marked for each voltage range unless the maximum and minimum values do not differ by more than 20 % of the mean value. The marked value shall not be less than 90 % of the maximum value.

d)	Equipment which an OPERATOR can set for different RATED supply voltages shall be provided with means for the indication of the voltage for which the equipment is set. For PORTABLE EQUIPMENT the indication shall be visible from the exterior. If the equipment is so constructed that the supply voltage setting can be altered without the use of a TOOL, the action of changing the setting shall also change the indication.

e)	Accessory MAINS socket-outlets accepting standard mains plugs shall be marked with the voltage if it is different from the MAINS supply voltage. If the outlet is for use only with specific equipment, it shall be marked to identify the equipment for which it is intended. If not, the maximum RATED current or power shall be marked, or symbol 13 of Table 1 placed beside the outlet with the full details included in the documentation.

### 5.1.4	Fuses

For any fuse which may be replaced by an OPERATOR, there shall be a marking beside the fuseholder, which will enable the OPERATOR to identify the correct replacement fuse.

### 5.1.5	TERMINALS, connections and operating devices

#### 5.1.5.1	General rule

If necessary for safety, an indication shall be given of the purpose of TERMINALS, connectors, controls, and indicators, including any connections for fluids such as gas, water and drainage. Where there is insufficient space, symbol 13 of Table 1 may be used.

Push-buttons and actuators of emergency stop devices, and indicators used only to indicate a warning of danger or the need for urgent action, shall be coloured red and coded as specified in GB/T 4025. If the meaning of colour relates to the safety of persons or the environment, supplementary means of coding shall be provided (per GB/T 4025).

#### 5.1.5.2	TERMINALS

TERMINALS for connection to the MAINS supply shall be identifiable. The following TERMINALS shall be marked as follows:

a)	FUNCTIONAL EARTH TERMINALS with Symbol 5 in Table 1.

b)	PROTECTIVE CONDUCTOR TERMINALS with symbol 6 of table 1, except when the PROTECTIVE CONDUCTOR TERMINAL is part of an approved MAINS appliance inlet. The symbol shall be placed on the TERMINAL, or close to it.

c)	TERMINALS of circuits to be connected to ACCESSIBLE conductive parts, with symbol 7 of Table 1 unless connection is self-evident.

d)	TERMINALS supplied from the interior of the equipment and which are HAZARDOUS LIVE, with the voltage, current, charge or energy value or range, or with symbol 13 of table 1. This requirement does not apply to MAINS supply outlets where a standard MAINS socket outlet is used;

### 5.1.6	Switches and circuit-breakers

If the power supply switch or circuit-breaker is used as the disconnecting device, the off-position shall be clearly marked.

Symbols 8 and 9 of Table 1 can, in some cases, also be as the device identification. A lamp alone is not considered to be a satisfactory marking.

If a push-button switch is used as the power supply switch, symbols 8 and 14 of Table 1 may be used to indicate the on-position, or symbols 9 and 15 to indicate the off-position, with the pair of symbols (8 and 9, or 9 and 15) close together.

### 5.1.7	Equipment protected by DOUBLE INSULATION or REINFORCED INSULATION

Equipment which is only partially protected by double insulation or reinforced insulation shall not be marked with symbol 10 of Table 1.

### 5.1.8	Field-wiring TERMINAL boxes

If the temperature of the TERMINALS or the ENCLOSURE of a field-wiring TERMINAL box or compartment exceeds 60 °C in NORMAL CONDITION at an ambient temperature of 40 °C, or at the maximum RATED ambient temperature of higher than 40 °C, there shall be a marking to warn the installer to consult the installation instructions before determining the temperature RATING of the cable to be connected to the TERMINALS. The marking shall be visible before and during connection, or be beside the TERMINALS. Symbol 13 of Table 1 is an acceptable marking.

Table 1 - Symbols

| No. | Symbol | Reference | Description |
| --- | --- | --- | --- |
| 1 |  | GB/T5465.2-5031 (2002-10) | Direct current |
| 2 |  | GB/T5465.2-5032 (2002-10) | Alternating current |
| 3 |  | GB/T5465.2-5033 (2002-10) | Both direct and alternating current |
| 4 |  | GB/T5465.2-5032 (2002-10) | Three-phase alternating current |
| 5 |  | GB/T5465.2-5017 (2006-08) | Earth (ground) terminal |
| 6 |  | GB/T5465.2-5019 (2006-08) | PROTECTIVE CONDUCTOR TERMINAL |
| 7 |  | GB/T5465.2-5020 (2002-10) | Frame or chassis TERMINAL |
| 8 |  | GB/T5465.2-5007 (2009-02) | On (power supply) |
| 9 |  | GB/T5465.2-5008 (2009-02) | Off (power supply) |
| 10 |  | GB/T5465.2-5172 (2003-02) | Equipment protected throughout by DOUBLE INSULATION or REINFORCED INSULATION |
| 11 |  |  | Caution, risk of electric shock |
| 12 |  | GB/T5465.2-5041 (2002-10) | Caution, hot surface |
| 13 |  | ISO7000-0434B (2004-01) | Caution |
| 14 |  | GB/T5465.2-5268 (2002-10) | In position of a bi-stable push control |
| 15 |  | GB/T5465.2-5269 (2002-10) | Out position of a bi-stable push control |
| 16 |  | ISO361 | Ionizing radiation |

## 5.2	Warning markings

Warning markings shall be visible when the equipment is ready for normal use. If a warning applies to a particular part of the equipment, the marking shall be placed on or near that part.

The size of warning markings shall be as follows:

a)	Symbols shall be at least 2,75 mm high. Text shall be at least 1,5 mm high and contrast in colour with the background.

b)	Symbols or text moulded, stamped or engraved in a material shall be at least 2,0 mm high. If not contrasting in colour, they shall have a depth or raised height of at least 0,5 mm.

If it is necessary for the responsible body or operator to refer to the documentation to preserve the protection afforded by the equipment, the equipment shall be marked with symbol 13 of Table 1. Symbol 13 of Table 1 is not required to be used if other symbols which address the safety aspects are used and explained in the documentation.

If the instructions for use state that an operator is permitted to gain access, using a tool, to a part or location which is in normal use may present a HAZARD, there shall be a warning marking indicating that the equipment shall be placed in a safe state before access.

## 5.3	Durability of markings

Required markings shall remain clear and legible under conditions of NORMAL USE and shall resist the effects of the cleaning agents specified by the manufacturer.

## 5.4	Files

### 5.4.1	General rule

The following documentation necessary for safety purposes, as needed by the operator or the responsible body, shall be provided with the equipment. Relevant safety documentation for service personnel authorized by the manufacturer shall be made available to such personnel:

a)	Intended use of the equipment;

b)	Technical specifications;

c)	Name and address of the manufacturer or supplier from whom technical assistance may be obtained;

d)	Information specified in 5.4.2 to 5.4.6;

e)	For equipment which for safety reasons requires specific accessories (for example probe assemblies) with specific characteristics, the documentation shall indicate that only accessories which meet the manufacturer’s specifications shall be used;

f)	If a hazard could be caused by an incorrect reading when measuring, indicating or detecting harmful or corrosive substances, or hazardous live electrical quantities, the instructions shall provide guidance on how to determine that the equipment is functioning correctly;

g)	Instructions for lifting and carrying.

Warning symbols and warning statements that are marked on the equipment shall be explained in the documentation. In particular, the documentation shall include a statement that it must be consulted in all cases where symbol 13 of Table 1 is marked, in order to find out the nature of the potential hazards and any actions which shall be taken to avoid them.

Documentation may be provided on printed or electronic media, however printed information shall be provided for all information necessary for safety that might not be available in electronic form at the time it is needed. The documentation shall be delivered with the equipment.

### 5.4.2	Equipment ratings

The documentation shall include the following:

a)	The supply voltage or voltage range, frequency or frequency range, and power or current rating;

b)	Description of all input and output connections;

c)	A statement on the range of environmental conditions for which the equipment is designed;

d)	A statement of the degree of ingress protection (IP), if the equipment is RATED according to GB/T 4208.

### 5.4.3	Equipment installation

The documentation shall include installation and specific commissioning instructions and, if necessary for safety, warnings against hazards which could arise during installation or commissioning or as a result of improper installation or commissioning of the equipment. Such information, if applicable, includes:

a)	Assembly, location and mounting requirements;

b)	Protective grounding instructions;

c)	Connection with a power supply;

d)	For permanently connected devices:

1)	Power wiring requirements;

2)	Requirements for any external switch or circuit-breaker and external overcurrent protection devices, and a recommendation that the switch or circuit-breaker be near the permanently connected equipment;

e)	Ventilation requirements;

f)	Requirements and safety characteristics for special external services;

g)	Instructions relating to sound level

### 5.4.4	Equipment operation

If applicable, Instructions for Use shall include:

a)	Identification and description of operating controls and their use in all operating modes;

b)	Instruction not to position the equipment so that it is difficult to operate the disconnecting device;

c)	Instructions for interconnection to accessories and other equipment, including indication of suitable accessories, detachable parts and any special materials;

d)	Specification of intermittent operation limits

e)	Explanations of symbols related to safety which are used on the equipment;

f)	Instructions for replacement of consumable materials;

g)	Instructions for cleaning and decontamination;

h)	A statement listing potentially poisonous or injurious substances that can be liberated from the equipment, and possible quantities;

i)	Detailed instructions about risk reduction procedures relating to flammable liquids;

j)	Details of methods of reducing the RISKS of burns from surfaces permitted to exceed the temperature limits of 10.1.

There shall be a statement in the instructions that, if the equipment is used in a manner not specified by the manufacturer, the protection provided by the equipment may be impaired.

### 5.4.5	Equipment maintenance and service

Instructions shall be provided to the responsible body in sufficient detail to permit safe maintenance, inspection and testing of the equipment, and to ensure continued safety of the equipment after the maintenance, inspection and test procedure.

If applicable, manufacturer’s documentation shall instruct against replacing detachable MAINS supply cords by inadequately rated cords.

For equipment using replaceable batteries, the specific battery type shall be stated.

The manufacturer shall specify any parts which are required to be examined or supplied only by the manufacturer or his agent.

The RATING and characteristics of replaceable fuses shall be stated.

Instructions on the following subjects shall be provided for service personnel, as necessary to permit safe servicing and continued safety of the equipment after servicing if the equipment is suitable to be serviced:

a)	Product-specific risks that may affect the safety of maintenance personnel;

b)	Protective measures against these risks;

c)	Verification of the safe state of the equipment after repair.

### 5.4.6	Integration into systems or effects resulting from special conditions

Aspects resulting from integration into systems or effects resulting from special ambient or application conditions shall be described in the documentation.

# 6	Protection against electric shock

## 6.1	General rule

Protection against electric shock shall be maintained in NORMAL CONDITION and SINGLE FAULT CONDITION. ACCESSIBLE parts shall not be hazardous live. Voltage, current, charge or energy between an accessible part and earth, or between any two accessible parts on the same piece of equipment within a distance of 1,8 m (over a surface or through air), shall not exceed the levels of 6.3.1 in normal condition nor of 6.3.2 in SINGLE FAULT CONDITION.

Accessible parts shall be prevented from becoming hazardous live in single fault condition. The primary means of protection (see 6.4) shall be supplemented by one of a), b), c) or d). Alternatively one of the single means of protection e) or f) shall be used:

a)	Protective bonding (see 6.5.1);

b)	Supplementary insulation (see 6.5.2);

c)	Automatic disconnection of the supply (see 6.5.5);

d)	current- or voltage-limiting device (see 6.5.6);

e)	Reinforced insulation (see 6.5.3);

f)	Protective impedance (see 6.5.4).

BASIC INSULATION

ENCLOSURES or BARRIER

Impedance

SUPPLEMENTARY INSULATION

PROTECTIVE BONDING

Automatic disconnection of the supply

Current- or voltage-limiting device

REINFORCED INSULATION

Protective impedance

HAZARDOUS part

DOUBLE INSULATION

Accessible part

and

Figure 1 - Acceptable arrangement of protective means against electric shock

## 6.2	Exceptions

If it is not feasible for operating reasons to prevent the following parts being both ACCESSIBLE and HAZARDOUS LIVE, they are permitted to be ACCESSIBLE to an OPERATOR during NORMAL USE while they are HAZARDOUS LIVE:

a)	parts of lamps and lamp sockets after lamp removal;

b)	parts intended to be replaced by an OPERATOR (for example, batteries) and which may be HAZARDOUS LIVE during the replacement or other OPERATOR, but only if they are ACCESSIBLE only by means of a TOOL and have a warning mark.

If any of the parts in a) and b) receives a charge from an internal capacitor, they shall not be HAZARDOUS LIVE 10 s after interruption of the supply.

## 6.3	Limit values for ACCESSIBLE parts

### 6.3.1	Limit values in normal conditions

Voltages above the levels of a) are deemed to be HAZARDOUS LIVE if any of the levels of b) or c) are exceeded at the same time.

a)	The a.c. voltage levels are 30 V r.m.s., 42,4 V peak and the d.c. voltage level is 60 V. For equipment intended for use in wet locations, the a.c. voltage levels are 16 V r.m.s., 22,6 V peak and the d.c.voltage level is 35 V.

b)	Current limits are:

1)	0,5 mA r.m.s. for sinusoidal waveforms, 0,7 mA peak for non-sinusoidal waveforms or mixed frequencies, or 2 mA d.c.

2)	70 mA r.m.s. when measured with the measuring circuit. This relates to possible burns at higher frequencies.

c)	The limits of capacitive charge or energy are:

1)	45μC charge for voltages up to peak 15kV or DC 15kV;

2)	350 mJ stored energy for voltages above 15 kV peak or d.c.

### 6.3.2	Levels in single fault condition

Voltages above the levels of a) are deemed to be HAZARDOUS LIVE if any of the levels of b) or c) are exceeded at the same time.

a)	The a.c. voltage levels are 50 V r.m.s., 70 V peak and the d.c. voltage level is 120 V. For equipment intended for use in wet locations, the a.c. voltage levels are 33 V r.m.s, 46,7 V peak and the d.c. voltage level is 70 V.

b)	Current limits are:

1)	3,5 mA r.m.s. for sinusoidal waveforms, 5 mA peak for non-sinusoidal waveforms or mixed frequencies, or 15 mA d.c.;

2)	500 mA r.m.s. when measured with the measuring circuit. This relates to possible burns at higher frequencies.

c)	The capacitance level is line B of Figure 2.

Key:

A--Normal conditions;

B--Single fault condition.

Figure 2 – Capacitance level versus voltage in NORMAL CONDITION and SINGLE FAULT CONDITION

## 6.4	Primary means of protection

### 6.4.1	ENCLOSURES and PROTECTIVE BARRIERS

If ENCLOSURES or PROTECTIVE BARRIERS provide protection by insulation, they shall meet the requirements of BASIC INSULATION.

If ENCLOSURES or PROTECTIVE BARRIERs provide protection by limiting access, CLEARANCES and CREEPAGE DISTANCES between ACCESSIBLE parts and HAZARDOUS LIVE parts shall meet the requirements of 6.6 and the applicable requirements for BASIC INSULATION.

### 6.4.2	BASIC INSULATION

CLEARANCES, CREEPAGE DISTANCES and solid insulation forming BASIC INSULATION between ACCESSIBLE PARTS and HAZARDOUS LIVE parts shall meet the requirements of 6.6.

### 6.4.3	Impedance

An impedance used as a primary means of protection meet all the following requirements:

a)	It shall limit the current or voltage to not more than the applicable level of 6.3.2;

b)	It shall be rated for the maximum working voltage and for the amount of power that it will dissipate;

c)	CLEARANCE and CREEPAGE DISTANCE between terminations of the impedance shall meet the applicable requirements of 6.6 for BASIC INSULATION.

## 6.5	Additional means of protection in case of SINGLE FAULT CONDITION

### 6.5.1	PROTECTIVE BONDING

#### 6.5.1.1	General rule

ACCESSIBLE conductive parts shall be bonded to the PROTECTIVE CONDUCTOR TERMINAL if they could become HAZARDOUS LIVE in case of a single fault of the primary means of protection specified in 6.4. Alternatively, such ACCESSIBLE parts shall be separated from parts which are HAZARDOUS LIVE by a conductive protective screen bonded to the PROTECTIVE CONDUCTOR TERMINAL.

#### 6.5.1.2	Integrity of PROTECTIVE BONDING

The integrity of the PROTECTIVE BONDING shall be assured as specified below:

a)	PROTECTIVE BONDING shall consist of direct connected structural parts or discrete conductors, or both.

b)	Soldered connections subject to mechanical stress shall be mechanically secured independently from the soldering. Such connections shall not be used for other purposes such as fixing constructional parts;

c)	Screw connections shall be secured against loosening.

d)	If a part of the equipment is removable by the OPERATOR, the PROTECTIVE BONDING for the remainder of the equipment shall not be interrupted;

e)	Unless they are specifically designed for electrical inter-connection and meet the requirements of movable conductive connections.

f)	The exterior metal braid of cables shall not be regarded as PROTECTIVE BONDING, even if connected to the PROTECTIVE CONDUCTOR TERMINAL.

g)	If power from the MAINS supply is passed through equipment for use by other equipment, means shall also be provided for passing the protective conductor through the equipment to protect the other equipment. The impedance of the protective conductor path through the equipment shall not exceed that specified in 6.5.1.4.

h)	Protective conductors may be bare or insulated. Insulation shall be green-and-yellow, except in the following cases:

1)	For earthing braids, either green-and-yellow or colourless-transparent;

2)	For internal protective conductors, and other conductors connected to the protective conductor terminal in assemblies, any colour may be used provided that no hazard is likely to arise from non-identification of the protective conductor.

Equipment using PROTECTIVE BONDING shall be provided with a TERMINAL that is suitable for connection to a protective conductor and meets the requirements of 6.5.1.3.

#### 6.5.1.3	PROTECTIVE CONDUCTOR TERMINAL

PROTECTIVE CONDUCTOR TERMINALS shall meet the following requirements:

a)	The contact surfaces shall be metal.

b)	The integral protective conductor connection of an appliance inlet shall be regarded as the PROTECTIVE CONDUCTOR TERMINAL.

c)	For equipment provided with a rewirable flexible cord and for PERMANENTLY CONNECTED EQUIPMENT, the PROTECTIVE CONDUCTOR TERMINAL shall be located near the MAINS supply TERMINALS.

d)	If the equipment does not require connection to a MAINS supply, but nevertheless has a circuit or part which is required to be protectively earthed, the PROTECTIVE CONDUCTOR TERMINAL shall be located near the TERMINALS of that circuit for which protective earthing is necessary. If this circuit has external TERMINALS, the PROTECTIVE CONDUCTOR TERMINAL shall also be external.

e)	PROTECTIVE CONDUCTOR TERMINALS for MAINS CIRCUITS shall be at least equivalent in current-carrying capacity to the MAINS supply TERMINALS.

f)	Plug-in type PROTECTIVE CONDUCTOR TERMINALS combined with other TERMINALS and intended to be connected and disconnected without the use of a TOOL, shall be designed so that the protective conductor connection makes first and breaks last with respect to the other connections.

g)	If the PROTECTIVE CONDUCTOR TERMINAL is also used for other bonding purposes, the protective conductor shall be applied first and secured independently of other connections. The protective conductor shall be connected in such a way that it is unlikely to be removed during servicing that does not require disconnection of the protective conductor.

h)	For equipment in which the protective conductor is required for protection against a single fault in a measuring circuit, the following shall apply:

1)	the PROTECTIVE CONDUCTOR TERMINAL and protective conductor shall have at least the current RATING of the measuring TERMINALS;

2)	the PROTECTIVE BONDING shall not be interrupted by any switching or interrupting device.

i)	FUNCTIONAL EARTH TERMINALS (for example, measuring earth TERMINALS) shall allow a connection which is independent from the connection of the protective conductor.

j)	If the PROTECTIVE CONDUCTOR TERMINAL is a binding screw assembly (see Figure 3), it shall be of a suitable size for the bond wire, but with a thread size no smaller than 4,0 mm, with at least three turns of the screw engaged.

k)	The contact pressure required for a bonding connection shall not be capable of being reduced by deformation of materials forming part of the connection.

Key:

A-Fixed part;

B-Gasket or splint;

C-Anti-stretching device;

D-Conductor space.

Figure 3 - Examples of binding screw assemblies

#### 6.5.1.4	Impedance of PROTECTIVE BONDING of plug-connected equipment

The impedance between the PROTECTIVE CONDUCTOR TERMINAL and each ACCESSIBLE part for which PROTECTIVE BONDING is specified shall not exceed 0,1 Ω. If the equipment has a non-detachable power cord, the impedance between the protective conductor plug pin of the MAINS cord and each ACCESSIBLE part for which PROTECTIVE BONDING is specified shall not exceed 0,2 Ω.

#### 6.5.1.5	Impedance of PROTECTIVE BONDING of PERMANENTLY CONNECTED EQUIPMENT

PROTECTIVE BONDING of PERMANENTLY CONNECTED EQUIPMENT shall be of low impedance.

If the equipment contains overcurrent protection devices for all poles of the MAINS supply, and if the wiring on the supply side of the overcurrent protection devices cannot become connected to ACCESSIBLE conductive parts in the case of a single fault, the test current need not be more than twice the RATED current of the internal overcurrent protection devices.

#### 6.5.1.6	Transformer PROTECTIVE BONDING screen

If the transformer is fitted with a screen for PROTECTIVE BONDING purposes that is separated only by BASIC INSULATION from a winding that is connected to a HAZARDOUS LIVE circuit, the screen shall satisfy the requirements of 6.5.1.2 a) and b), and be of low impedance.

### 6.5.2	SUPPLEMENTARY INSULATION

CLEARANCES, CREEPAGE DISTANCES and solid insulation forming SUPPLEMENTARY INSULATION shall meet the applicable requirements of 6.6.

### 6.5.3	REINFORCED INSULATION

CLEARANCES, CREEPAGE DISTANCES and solid insulation forming REINFORCED INSULATION shall meet the applicable requirements of 6.6.

### 6.5.4	PROTECTIVE IMPEDANCE

A PROTECTIVE IMPEDANCE shall limit the current or voltage to the levels of 6.3.1 in NORMAL CONDITION and 6.3.2 in single fault condition.

Insulation between the terminations of the PROTECTIVE IMPEDANCE shall meet the requirements of 6.6 for DOUBLE INSULATION or REINFORCED INSULATION. A PROTECTIVE IMPEDANCE shall be one or more of the following:

a)	An appropriate single component which shall be constructed, selected and tested so that safety and reliability for protection against electric shock is assured. In particular, the component shall be:

1)	rated for twice the maximum working voltage;

2)	if a resistor, rated for twice the power dissipation for the maximum working voltage.

b)	It shall be a combination of components.

A protective impedance shall not be a single electronic device that employs electron conduction in a vacuum, gas or semiconductor

### 6.5.5	Automatic disconnection of the supply

An automatic disconnection device shall meet both the following requirements:

a)	Rated to disconnect the load within specified time

b)	It shall be RATED for the maximum RATED load conditions of the equipment.

### 6.5.6	Current- or voltage-limiting device

A current- or voltage-limiting device shall meet all the following requirements:

a)	It shall be rated to limit the current or voltage to the levels not exceeding the values of 6.3.2;

b)	It shall be rated for the maximum working voltage and, if applicable, for the maximum operational current;

c)	CLEARANCE and CREEPAGE DISTANCE between the terminations of the current or voltage-limiting device shall meet the applicable requirements of 6.6 for SUPPLEMENTARY INSULATION.

## 6.6	Insulation requirements

### 6.6.1	The nature of insulation

#### 6.6.1.1	General rule

Insulation between circuits and ACCESSIBLE parts or between separate circuits consists of a combination of CLEARANCES, CREEPAGE DISTANCES and solid insulation. When used to provide protection against a HAZARD, the insulation needs to withstand the electric stresses that are caused by the voltages that may appear on the MAINS supply or in the equipment.

#### 6.6.1.2	CLEARANCES

If the equipment is RATED to operate at an altitude above 2 000 m, all CLEARANCES shall be multiplied by the applicable factor of Table 2.

Table 2 - Multiplication factors for CLEARANCES of equipment RATED for operation at altitudes up to 5 000 m

| RATED operating altitude m | Multiplication factor |
| --- | --- |
| ≤2000 | 1.00 |
| 2001~3000 | 1.14 |
| 3001~4000 | 1.29 |
| 4001~5000 | 1.48 |

#### 6.6.1.3	CREEPAGE DISTANCES

According to their comparative tracking index (CTI) values, the materials are divided into four groups as follows:

--Material group I	600≤CTI;

--Material group II	400≤CTI<600;

--Material group Ⅲa	175≤CTI<400;

--Material group Ⅲb	100≤CTI<175.

For materials where the CTI value is not known, material group IIIb is assumed.

For glass, ceramics, or other inorganic insulating materials which do not track, there are no requirements for CREEPAGE DISTANCES.

#### 6.6.1.4	Solid insulation

Solid insulation is not a renewable medium: damage is cumulative over the life of the equipment. Solid insulation is also subject to aging and to degradation from repeated high voltage testing.

#### 6.6.1.5	Requirements for insulation according to type of circuit

Requirements for insulation in particular types of circuits are specified as follows:

a)	In 6.6.2 for MAINS CIRCUITS of OVERVOLTAGE CATEGORY II with a nominal supply voltage up to 300 V.

b)	In 6.6.3 for secondary circuits separated from the circuits in a) only by means of a transformer.

c)	The mains circuit of overvoltage category III or IV or overvoltage category II exceeding 300V is in accordance with A.1 in Annex A.

d)	In A.2 for secondary circuits separated from the circuits in c) only by means of a transformer.

e)	Circuits that have one or more of the following characteristics is in accordance with A.3:

1)	the maximum possible TRANSIENT OVERVOLTAGE is limited by the supply source or within the equipment to a known level below the level assumed for the MAINS CIRCUIT;

2)	the maximum possible TRANSIENT OVERVOLTAGE is above the level assumed for the MAINS CIRCUIT;

3)	the WORKING VOLTAGE is the sum of voltages from more than one circuit, or is a mixed voltage;

4)	the WORKING VOLTAGE includes a recurring peak voltage that may include a periodic non-sinusoidal waveform or a non-periodic waveform that occurs with some regularity;

5)	the WORKING VOLTAGE has a frequency above 30 kHz.

### 6.6.2	Insulation for MAINS CIRCUITS of OVERVOLTAGE CATEGORY II with a nominal supply voltage up to 300 V

#### 6.6.2.1	CLEARANCES and CREEPAGE DISTANCES

CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS shall meet the values of Table 3 and take the following into account:

a)	The values in Table 3 are for BASIC INSULATION and SUPPLEMENTARY INSULATION. Values for REINFORCED INSULATION shall be twice the values for BASIC INSULATION.

b)	Minimum CLEARANCE for BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION for POLLUTION DEGREE 3 is 0,8 mm;

c)	If the equipment is RATED to operate at an altitude above 2 000 m, the CLEARANCES shall be multiplied by the applicable factor of Table 2.

Table 3 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOLTAGE CATEGORY II up to 300 V

| Voltage line-to-neutral a.c. r.m.s. or d.c. | Values for CLEARANCE | Values for CREEPAGE DISTANCE |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Printed wiring board material |  | Other insulating material |  |  |  |  |  |  |
|  |  | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 |  |  | POLLUTION DEGREE 3 |  |  |
|  |  | All material groups | Material groups Ⅰ, Ⅱ, Ⅲa | All material groups | Material group Ⅰ | Material group Ⅱ | Material group Ⅲa | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ |
| V | mm | mm | mm | mm | mm | mm | mm | mm | mm | mm |
| ≤150 | 0.5 | 0.5 | 0.5 | 0.5 | 0.8 | 1.1 | 1.6 | 2.0 | 2.2 | 2.5 |
| >150~300 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 2.1 | 3.0 | 3.8 | 4.1 | 4.7 |

Linear interpolation of the CREEPAGE DISTANCES is allowed.

#### 6.6.2.2	Solid insulation

##### 6.6.2.2.1	General rule

Solid insulation of MAINS CIRCUITS shall withstand the electric and mechanical stresses that may occur in NORMAL USE, in all RATED environmental conditions (see Clause 4), during the intended life of the equipment.

Solid insulation shall also meet the following requirements, as applicable:

a)	For solid insulation used as an ENCLOSURE or PROTECTIVE BARRIER, the requirements of Clause 8;

b)	For moulded and potted parts, the requirements of 6.6.2.2.2;

c)	For the inner insulating layers of printed wiring boards, the requirements of 6.6.2.2.3;

d)	For thin-film insulation, the requirements of 6.6.2.2.4.

##### 6.6.2.2.2	Moulded and potted parts

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located on an interface between the same two layers moulded together shall be separated by at least 0,4 mm (see Figure 4, item L) after the moulding is completed.

Key:

1--Layer 1;

2--Layer 2;

C--Conductor;

L--Conductor spacing.

Figure 4 - Distance between conductors on an interface between two layers

##### 6.6.2.2.3	Inner insulating layers of printed wiring boards

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers shall be separated by at least 0,4 mm (see Figure 5, item L).

Key:

L--Distance between adjacent conductors;

A--Layer;

C--Conductor.

Figure 5 - Distance between adjacent conductors along an interface of two inner layers

REINFORCED INSULATION of inner insulating layers of printed wiring boards shall also have adequate electric strength through the respective layers. One of the following methods shall be used:

a)	The thickness of the insulation is at least 0,4 mm;

b)	The insulation is assembled from at least two separate layers of printed wiring board materials, each of which is RATED by the manufacturer of the material for an electric strength of at least the value of the test voltage of Table 4 for BASIC INSULATION;

c)	The insulation is assembled from at least two separate layers of printed wiring board materials, and the combination of layers is RATED by the manufacturer of the material for an electric strength of at least the value of the test voltage of Table 4 for REINFORCED INSULATION.

Table 4 - Test voltages for solid insulation in mains circuits of OVERVOLTAGE CATEGORY II Up to 300 V

| Voltage line-to-neutral a.c. r.m.s. or d.c. | 1 min a.c. test voltage |  | 1 min d.c. test voltage |  |
| --- | --- | --- | --- | --- |
|  | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION |
| V | V | V | V | V |
| ≤150 | 1350 | 2700 | 1900 | 3800 |
| >150~300 | 1500 | 3000 | 2100 | 4200 |

##### 6.6.2.2.4	Thin-film insulation

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers (see Figure 6, item L).shall be separated by the applicable CLEARANCE and CREEPAGE DISTANCE of 6.6.2.1.

Key:

L--Distance between adjacent conductors;

A--Layers of thin-film material, such as tape and polyester film;

C--Conductor.

NOTE: There may be air present between the layers.

Figure 6 - Distance between adjacent conductors located between the same two layers

REINFORCED INSULATION through the layers of thin-film insulation shall also have adequate electric strength. One of the following methods can be used.

a)	The thickness of the insulation is at least 0,4 mm;

b)	The insulation consists of at least two separate layers of thin-film materials, each of which is RATED by the manufacturer of the material for an electric strength of at least the value of the test voltage of Table 4 for BASIC INSULATION.

c)	The insulation consists of at least three separate layers of thin-film materials, any two of which have been tested to exhibit adequate electric strength.

### 6.6.3	Insulation for secondary circuits derived from MAINS CIRCUITS of OVERVOLTAGE CATEGORY II up to 300 V

#### 6.6.3.1	General rule

In this standard, the isolation of the secondary circuits and the mains circuits is realized by a transformer, in which the primary windings are separated from the secondary windings by reinforced insulation, double insulation, or a screen connected to the protective terminal.

#### 6.6.3.2	CLEARANCES

CLEARANCES for secondary circuits shall meet a) or b).

a)	For BASIC INSULATION and SUPPLEMENTARY INSULATION, meet the values of Table 5, or for REINFORCED INSULATION meet twice the values of Table 5.

b)	Pass the voltage test using the applicable test voltage from Table 6. The following adjustments apply:

1)	values for test voltages for REINFORCED INSULATION are 1,6 times the values for BASIC INSULATION;

2)	if the equipment is RATED to operate at an altitude above 2 000 m, the values for CLEARANCES are multiplied by the applicable factor of Table 2;

3)	minimum CLEARANCE is 0,2 mm for POLLUTION DEGREE 3 and 0,8 mm for POLLUTION DEGREE 3.

Table 5 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOLTAGE CATEGORY II up to 300 V

| Secondary Working voltage |  | Mains voltage, line-to-neutral, overvoltage category II |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | ≤150V a.c. r.m.s. |  | >150V~300V a.c. r.m.s. |  |
| a.c. r.m.s. V | d.c. or a.c. peak V | CLEARANCES mm | Test Voltage (AC RMS) V | CLEARANCES mm | Test Voltage (AC RMS) V |
| 16 | 22.6 | 0.10 | 500 | 0.48 | 830 |
| 30 | 42.4 | 0.11 | 510 | 0.50 | 840 |
| 50 | 70 | 0.12 | 520 | 0.53 | 860 |
| 100 | 140 | 0.13 | 540 | 0.61 | 900 |
| 150 | 210 | 0.16 | 580 | 0.69 | 940 |
| 300 | 420 | 0.39 | 770 | 0.94 | 1040 |
| 600 | 840 | 1.01 | 1070 | 1.61 | 1450 |
| 1000 | 1400 | 1.92 | 1630 | 2.52 | 1970 |
| 1250 | 1750 | 2.50 | 1960 | 3.16 | 2280 |
| 1600 | 2240 | 3.39 | 2390 | 4.11 | 2730 |
| 2000 | 2800 | 4.49 | 2890 | 5.30 | 3230 |
| 2500 | 3500 | 6.02 | 3520 | 6.91 | 3850 |
| 3200 | 4480 | 8.37 | 4390 | 9.16 | 4660 |
| 4000 | 5600 | 10.9 | 5320 | 11.6 | 5610 |
| 5000 | 7000 | 14.0 | 6590 | 14.9 | 6960 |
| 6300 | 8820 | 18.2 | 8270 | 19.1 | 8620 |
| 8000 | 11200 | 23.9 | 10400 | 24.7 | 10700 |
| 10000 | 14000 | 30.7 | 12900 | 31.6 | 13300 |
| 12500 | 17500 | 39.6 | 16100 | 40.5 | 16400 |
| 16000 | 22400 | 52.5 | 20400 | 53.5 | 20700 |
| 20000 | 28000 | 67.9 | 25300 | 68.9 | 25600 |
| 25000 | 35000 | 87.9 | 31600 | 89.0 | 32000 |
| 32000 | 44800 | 117 | 40400 | 118 | 40700 |
| 40000 | 56000 | 151 | 50300 | 153 | 50800 |
| 50000 | 70000 | 196 | 62800 | 198 | 63400 |
| 63000 | 88200 | 258 | 79400 | 260 | 80000 |
| Linear interpolation is allowed |  |  |  |  |  |

#### 6.6.3.3	CREEPAGE DISTANCES

CREEPAGE DISTANCES for BASIC INSULATION or SUPPLEMENTARY INSULATION for secondary circuits shall meet the applicable values of Table 6, based on the WORKING VOLTAGE which stresses the insulation. Values for REINFORCED INSULATION are twice the values for BASIC INSULATION.

Table 6 - CREEPAGE DISTANCES for secondary circuits

| Secondary WORKING VOLTAGE a.c. r.m.s. or d.c. | Printed wiring board material |  | Other insulating material |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 |  |  | POLLUTION DEGREE 3 |  |  |
|  | All material groups | Material groups I, II, or IIIa | All material groups | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ | Material group Ⅰ | Material group Ⅱ | Material group Ⅲb |
| V | mm | mm | mm | mm | mm | mm | mm | mm | mm |
| 10 | 0.025 | 0.04 | 0.08 | 0.40 | 0.40 | 0.40 | 1.00 | 1.00 | 1.00 |
| 12.5 | 0.025 | 0.04 | 0.09 | 0.42 | 0.42 | 0.42 | 1.05 | 1.05 | 1.05 |
| 16 | 0.025 | 0.04 | 0.10 | 0.45 | 0.45 | 0.45 | 1.10 | 1.10 | 1.10 |
| 20 | 0.025 | 0.04 | 0.11 | 0.48 | 0.48 | 0.48 | 1.20 | 1.20 | 1.20 |
| 25 | 0.025 | 0.04 | 0.125 | 0.50 | 0.50 | 0.50 | 1.25 | 1.25 | 1.25 |
| 32 | 0.025 | 0.04 | 0.14 | 0.53 | 0.53 | 0.53 | 1.3 | 1.3 | 1.3 |
| 40 | 0.025 | 0.04 | 0.16 | 0.56 | 0.80 | 1.10 | 1.4 | 1.6 | 1.8 |
| 50 | 0.025 | 0.04 | 0.18 | 0.60 | 0.85 | 1.20 | 1.5 | 1.7 | 1.9 |
| 63 | 0.040 | 0.063 | 0.20 | 0.63 | 0.90 | 1.25 | 1.6 | 1.8 | 2.0 |
| 80 | 0.063 | 0.10 | 0.22 | 0.67 | 0.95 | 1.3 | 1.7 | 1.9 | 2.1 |
| 100 | 0.10 | 0.16 | 0.25 | 0.71 | 1.00 | 1.4 | 1.8 | 2.0 | 2.2 |
| 125 | 0.16 | 0.25 | 0.28 | 0.75 | 1.05 | 1.5 | 1.9 | 2.1 | 2.4 |
| 160 | 0.25 | 0.40 | 0.32 | 0.80 | 1.1 | 1.6 | 2.0 | 2.2 | 2.5 |
| 200 | 0.40 | 0.63 | 0.42 | 1.00 | 1.4 | 2.0 | 2.5 | 2.8 | 3.2 |
| 250 | 0.56 | 1.0 | 0.56 | 1.25 | 1.8 | 2.5 | 3.2 | 3.6 | 4.0 |
| 320 | 0.75 | 1.6 | 0.75 | 1.60 | 2.2 | 3.2 | 4.0 | 4.5 | 5.0 |
| 400 | 1.0 | 2.0 | 1.0 | 2.0 | 2.8 | 4.0 | 5.0 | 5.6 | 6.3 |
| 500 | 1.3 | 2.5 | 1.3 | 2.5 | 3.6 | 5.0 | 6.3 | 7.1 | 8.0 |
| 630 | 1.8 | 3.2 | 1.8 | 3.2 | 4.5 | 6.3 | 8.0 | 9.0 | 10.0 |
| 800 | 2.4 | 4.0 | 2.4 | 4.0 | 5.6 | 8.0 | 10.0 | 11 | 12.5 |
| 1000 | 3.2a | 5.0a | 3.2 | 5.0 | 7.1 | 10.0 | 12.5 | 14 | 16 |
| 1250 |  |  | 4.2 | 6.3 | 9,0 | 12.5 | 16 | 18 | 20 |
| 1600 |  |  | 5.6 | 8.0 | 11 | 16 | 20 | 22 | 25 |
| 2000 |  |  | 7.5 | 10.0 | 14 | 20 | 25 | 28 | 32 |
| 2500 |  |  | 10.0 | 12.5 | 18 | 25 | 32 | 36 | 40 |
| 3200 |  |  | 12.5 | 16 | 22 | 32 | 40 | 45 | 50 |
| 4000 |  |  | 16 | 20 | 28 | 40 | 50 | 56 | 63 |
| 5000 |  |  | 20 | 25 | 36 | 50 | 63 | 71 | 80 |
| 6300 |  |  | 25 | 32 | 45 | 63 | 80 | 90 | 100 |
| 8000 |  |  | 32 | 40 | 56 | 80 | 100 | 110 | 125 |
| 10000 |  |  | 40 | 50 | 71 | 100 | 125 | 140 | 160 |
| 12500 |  |  | 50 | 63 | 90 | 125 |  |  |  |
| 16000 |  |  | 63 | 80 | 110 | 160 |  |  |  |
| 20000 |  |  | 80 | 100 | 140 | 200 |  |  |  |
| 25000 |  |  | 100 | 125 | 180 | 250 |  |  |  |
| 32000 |  |  | 125 | 160 | 220 | 320 |  |  |  |
| 40000 |  |  | 160 | 200 | 280 | 400 |  |  |  |
| 50000 |  |  | 200 | 250 | 360 | 500 |  |  |  |
| 63000 |  |  | 250 | 320 | 450 | 600 |  |  |  |
| Linear interpolation is allowed |  |  |  |  |  |  |  |  |  |
| aFor voltages above 1 000 V, CREEPAGE DISTANCES on printed wiring board material are the same as for other insulators of the same Material Group. bMaterial group Ⅲb is not recommended for application in POLLUTION DEGREE 3 above 630 V. |  |  |  |  |  |  |  |  |  |

#### 6.6.3.4	Solid insulation

##### 6.6.3.4.1	General rule

Solid insulation in secondary circuits shall withstand the electrical and mechanical stresses that may occur in NORMAL USE, in all RATED environmental conditions (see Clause 4), during the intended life of the equipment.

Conformity is checked by both of the following tests:

a)	For solid insulation used as an ENCLOSURE or BARRIER, the requirements of Clause 8;

b)	For moulded and potted parts, the requirements of 6.6.3.4.2;

c)	For inner insulating layers of printed wiring boards, the requirements of 6.6.3.4.3;

d)	For thin-film insulation, the requirements of 6.6.3.4.4.

##### 6.6.3.4.2	Moulded and potted parts

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers shall be separated by the applicable minimum distance of Table 7 (see Figure 5, item L).

Table 7 - Minimum values for distance or thickness (see 6.6.3.4.2 to 6.6.3.4.4)

| AC peak or DC Operating voltage or repeated peak voltage kV | Minimum value mm | Peak value of the a.c. or d.c. WORKING VOLTAGE or recurring peak voltage kV | Minimum value mm |
| --- | --- | --- | --- |
| >0.0467~0.33 | 0.05 | >8.0~10 | 3.5 |
| >0.33~0.8 | 0.1 | >10~12 | 4.5 |
| >0.8~1.0 | 0.15 | >12~15 | 5.5 |
| >1.0~1.2 | 0.2 | >15~20 | 8 |
| >1.2~1.5 | 0.3 | >20~25 | 10 |
| >1.5~2.0 | 0.45 | >25~30 | 12.5 |
| >2.0~2.5 | 0.6 | >30~40 | 17 |
| >2.5~3.0 | 0.8 | >40~50 | 22 |
| >3.0~4.0 | 1.2 | >50~60 | 27 |
| >4.0~5.0 | 1.5 | >60~80 | 35 |
| >5.0~6.0 | 2 | >80~100 | 45 |
| >6.0~8.0 | 3 |  |  |

##### 6.6.3.4.3	Inner insulating layers of printed wiring boards

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers shall be separated by the applicable minimum distance of Table 7 (see Figure 5, item L).

REINFORCED INSULATION of inner insulating layers of printed wiring boards shall also have adequate electric strength through the respective layers. One of the following methods shall be used:

a)	The thickness of the insulation is at least the applicable minimum distance of Table 7;

b)	The insulation is assembled from at least two separate layers of printed wiring board materials, each of which is RATED by the manufacturer of the material for an electric strength at least the value of the test voltage of Table 5 for BASIC INSULATION;

c)	The insulation is assembled from at least two separate layers of printed wiring board materials, and the combination of layers is RATED by the manufacturer of the material for an electric strength at least the value of the test voltage of Table 5 multiplied by 1,6.

##### 6.6.3.4.4	Thin-film insulation

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers (see Figure 6, item L) shall be separated by at least the applicable CLEARANCE and CREEPAGE DISTANCE of 6.6.3.2 and 6.6.3.3.

REINFORCED INSULATION through the layers of thin-film insulation shall also have adequate electric strength. One of the following methods shall be used:

a)	The thickness of the insulation is at least the applicable value of Table 7;

b)	The insulation consists of at least two separate layers of thin-film materials, each of which is RATED by the manufacturer of the material for an electric strength of at least the value of the test voltage of Table 5 for BASIC INSULATION;

c)	The insulation consists of at least three separate layers of thin-film materials, any two of which have been tested to exhibit adequate electric strength.

# 7	Protection against mechanical HAZARDS

## 7.1	General rule

The equipment shall not cause a mechanical HAZARD in NORMAL USE, or cause a HAZARD in a SINGLE-FAULT CONDITION that might not be easily noticed. Mechanical hazards include, but are not limited to, the following:

a)	Sharp edges that may cause cuts (see 7.2);

b)	Moving parts that could crush body parts or penetrate the skin (see 7.3);

c)	Unstable equipment that could fall on a person while in use or while being moved (see 7.4);

d)	Falling equipment, resulting from breakage of the carrying device (see 7.5), wall mounting bracket (see 7.6) or other support part (see 7.5); and

e)	Expelled parts from the equipment (see 7.7).

## 7.2	Sharp edges

All easily-touched parts of the equipment shall be smooth and rounded so as not to cause injury during NORMAL USE of the equipment.

Unless the fault presents an obvious hazard, easily-touched parts of the equipment shall not cause an injury in single fault condition.

## 7.3	Moving parts

### 7.3.1	General rule

HAZARDS from moving parts shall not exceed a tolerable level except as specified in 7.3.2. The conditions specified in 7.3.4 and 7.3.5 are considered to represent a tolerable level.

### 7.3.2	Exceptions

If it is not feasible for operating reasons to prevent certain moving parts from causing a potential hazard, access is permitted in the following circumstances.

a)	Equipment with easily-touched moving parts which are obviously intended to operate on parts or materials external to the equipment, for example drilling and mixing equipment. Such equipment shall be designed to minimize inadvertent touching of these moving parts (for example, by guards or handles).

b)	If, during routine maintenance outside NORMAL USE, it is unavoidable for technical reasons for an OPERATOR to perform a function which requires access to moving parts that could cause a HAZARD, access is permitted if all of the following precautions have been taken.

1)	the contact of operating components should not be touched without tools;

2)	the instructions for the RESPONSIBLE BODY include a statement that OPERATORS shall be trained before being allowed to perform the hazardous operation;

3)	there are warning markings on covers or parts which shall be removed to obtain access, prohibiting access by untrained OPERATORS. As an alternative, symbol 13 of Table 1 shall be placed on the covers or parts, with the warnings included in the documentation.

### 7.3.3	RISK assessment for mechanical HAZARDS to body parts

Risks reduced to a tolerable level by at least applicable minimum protective measure of Table 12, taking into account the severity, probability of exposure and possibility of avoiding the hazard

Table 8 - Protective measures against mechanical HAZARDS to body parts

| Mechanical HAZARD conditions |  |  | Minimum protective measures d |
| --- | --- | --- | --- |
| Severity a | Probability of exposure b | Probability of avoiding the HAZARD c |  |
| S | E2 | P2 | C |
| S | E2 | P1 | C |
| S | E1 | P2 | C |
| S | E1 | P1 | B |
| M | E2 | P2 | B |
| M | E2 | P1 | A |
| M | E1 | P2 | A |
| M | E1 | P1 | 无 |
| a Severity: M--Moderate hazards, sufficient to bruise or scratch a body part; S - Serious hazards, sufficient to break bones or dismember a body part. b Possibility of exposure: E 1 - Exposure is not intended during normal use; E 2 - Exposure is intended during normal use. c Possibility to avoid danger: P 1 - Possibility of avoidance The movement is visible and the velocity is low enough to keep the body part away and not trapped; or activate a visual or audible alarm before a body part becomes trapped. P 2 —No possibility of avoidance Conditions other than P1. d Minimum protective measures: A—Low level measures: warning markings, audible or visual signals or instructions for use; B--Moderate measures; emergency switches, Protective barriers or covers removable only with a tool, distances (see ISO 13857) or separations (see ISO 13854 or EN 349); C—Low level measures: warning markings, audible or visual signStringent measures; interlocks, PROTECTIVE BARRIERS or covers removable only with a TOOL and with instructions to disconnect from the power source. |  |  |  |

### 7.3.4	Limitation of force and pressure

The levels below shall be met in NORMAL CONDITION and SINGLE-FAULT CONDITION.

The maximum tolerable continuous contact pressure is 50 N/cm², with a maximum force of 150 N.

The maximum tolerable temporary force for body contact areas of at least 3 cm 2 is 250 N for a maximum duration of 0.75 s.

### 7.3.5	Gap limitations between moving parts

#### 7.3.5.1	Gap limitations between moving parts – Access normally allowed

If the force and pressure of the moving parts exceed the limits of 7.3.4 and if a body part could be inserted between moving parts, the width of the gap shall not decrease from a value larger than the minimum gap of Table 9 for that body part to a value smaller than the minimum gap in NORMAL CONDITION and SINGLE-FAULT CONDITION.

#### 7.3.5.2	Gap limitations between moving parts – Access normally prevented

While parts are moving, gaps between moving parts into which body parts could be inserted shall not increase to more than the acceptable gap of Table 10, in NORMAL CONDITION and SINGLE-FAULT CONDITION.

Table 9 - Minimum maintained gaps to prevent crushing for different body parts

| Part of body | Minimum gap (a) to avoid crushing mm | Illustration |
| --- | --- | --- |
| Torso | 500 |  |
| Head | 300 |  |
| Leg | 180 |  |
| Foot | 120 |  |
| Toes | 50 |  |

Table 9 - Minimum maintained gaps to prevent crushing for different body parts(continued)

| Part of body | Minimum gap (a) to avoid crushing mm | Illustration |
| --- | --- | --- |
| Arm | 120 |  |
| Hand, wrist, fist | 100 |  |
| Finger | 25 |  |
| The values in this table are for adults. For equipment that may be operated by children or youths, consideration should be given to smaller dimensions |  |  |

Table 10 - Maximum gaps to prevent access for different body parts

| Part of body | Maximum gap (aa) to prevent access mm |
| --- | --- |
| Head | 120 |
| Foot | 35 |
| Finger | 4 |
| The values in this table are for adults. For equipment that may be operated by children or youths, consideration should be given to smaller dimensions |  |
| a See Table 9 for examples of gap (a). |  |

## 7.4	Stability

Equipment and assemblies of equipment not secured to the building structure before operation shall be physically stable.

If means are provided to ensure that stability is maintained after the opening of drawers, etc. by an OPERATOR, either these means shall be automatic or there shall be a warning marking to apply the means.

## 7.5	Provisions for lifting and carrying

### 7.5.1	General rule

Equipment or parts having a mass of 18 kg or more shall be provided with a means for lifting and carrying, or directions shall be given in the documentation.

### 7.5.2	Handles and grips

If carrying handles or grips are fitted to the equipment, or supplied with it, they shall be capable of withstanding a force of four times the weight of the equipment.

### 7.5.3	Lifting devices and supporting parts

Parts of lifting devices and parts that support heavy loads shall be RATED for the maximum load or shall be tested to withstand four times the maximum static load.

## 7.6	Wall mounting

Mounting brackets on equipment intended to be mounted on a wall or ceiling shall withstand a force of four times the weight of the equipment.

## 7.7	Expelled parts

Equipment shall contain or limit the energy of parts which could cause a hazard if expelled in the event of a fault. The means of protection against expelled parts shall not be removable without the aid of a tool.

# 8	RESISTANCE TO MECHANICAL STRESSES

Equipment shall not cause a hazard when subjected to mechanical stresses likely to occur in NORMAL USE. The normal energy protection level required is 5 J. Levels below 5 J but not less than 1 J are permitted provided that all the following criteria are met.

a)	The lower level is justified by a risk assessment carried out by the manufacturer.

b)	When the equipment is installed in its intended application, it cannot easily be touched by unauthorized persons or the general public.

c)	In normal use, the equipment is only accessed for occasional operations such as adjustment, programming or maintenance.

d)	The equipment is marked with an IK code in accordance with GB/T 20138 or with symbol 13 of Table 1, and the rated energy level and test method are stated in the accompanying documents. For non-metallic enclosures with a minimum rated ambient temperature below 2 °C, the stated value shall be that applicable to the lowest rated ambient temperature. If impact energies used are between IK values of GB/T 20138, any IK marking shall be for the nearest lower value.

# 9	Protection against the spread of fire

## 9.1	General rule

There shall be no spread of fire outside the equipment in NORMAL CONDITION or in SINGLE FAULT CONDITION. See Figure 7 for the flow chart of the requirements for protection against the spread of fire.

There shall be no spread of fire outside the equipment in NORMAL CONDITION or in SINGLE FAULT CONDITION

Testing in the SINGLE FAULT CONDITIONS that could cause the spread of fire outside the equipment

Eliminating or reducing the sources of ignition within the equipment

Energy limitation

Separation requirements

Requirement related to flammable liquids

Construction requirements for components

Enclosure requirements

Requirements for flammable liquids

Containment of the fire within the equipment, should it occur

Energizing of the equipment is controlled by a switch that is held closed by the OPERATOR

Testing of circuits designed to produce heat in SINGLE FAULT CONDITIONS which could cause ignition

Overcurrent protection

ACCEPTABLE

Figure 7 - Flow chart to explain the requirements for protection against the spread of fire

## 9.2	Eliminating or reducing the sources of ignition within the equipment

The possibility of ignition and occurrence of fire is considered to be reduced to a tolerable level if all the following requirements, if applicable, are met.

a)	1) or 2):

1)	the voltage, current and power available to the circuit or part of equipment are limited as specified in 9.4;

2)	insulation between parts at different potentials meets the requirements for BASIC INSULATION, or it can be demonstrated that bridging the insulation will not cause ignition.

b)	Any ignition hazard related to flammable liquids is reduced to a tolerable level as specified in 9.5.

c)	In circuits designed to produce heat, no ignition occurs when tested in SINGLE FAULT CONDITION.

## 9.3	Containment of the fire within the equipment, should it occur

### 9.3.1	General rule

The risk of the spread of fire outside the equipment is considered to be reduced to a tolerable level if the equipment meets one of the following constructional requirements.

a)	Energizing of the equipment is controlled by a switch that needs to be continuously held in the energized state by the OPERATOR.

b)	The equipment and the equipment ENCLOSURE conform to the constructional requirements of 9.3.2 and the applicable requirements of 9.5 are met.

### 9.3.2	Constructional requirements

The following constructional requirements shall be met.

a)	Connectors and insulating material on which components are mounted shall have a flammability classification V-2, or better, of GB/T 5169.16, see Clause 14.6 for the requirements of printed wiring board.

b)	Insulated wires and cables shall retard flame propagation.

c)	The ENCLOSURE shall meet the following requirements:

1)	The bottom and sides of the ENCLOSURE within the 5˚ sarc of Figure 9 of circuits that are not limited circuits according to 9.4 shall comply with one of the following requirements:

--No perforations;

--be made of metal with perforations as specified in Table 11;

--be a metal screen with a mesh not exceeding 2 mm × 2 mm centre to centre and a wire diameter of at least 0.45 mm;

--have openings with baffles according to Figure 8.

2)	The ENCLOSURE, and any baffle or flame barrier, shall be made of metal (except magnesium) or of non-metallic materials having a flammability classification of V-1 or better, of GB/T 5169.16.

3)	The ENCLOSURE, and any baffle or flame barrier, shall have adequate rigidity.

Table 11 – Acceptable perforation of the bottom of an ENCLOSURE

| Minimum thickness mm | Maximum diameter of holes mm | Minimum spacing of holes centre to centre mm |
| --- | --- | --- |
| 0.66 | 1.14 | 1.70 (233 holes/645 mm2) |
| 0.66 | 1.19 | 2.36 |
| 0.76 | 1.15 | 1.70 |
| 0.76 | 1.19 | 2.36 |
| 0.81 | 1.91 | 3.18 (72 holes / 645 mm 2) |
| 0.89 | 1.90 | 3.18 |
| 0.91 | 1.60 | 2.77 |
| 0.91 | 1.98 | 3.18 |
| 1.00 | 1.60 | 2.77 |
| 1.00 | 2.00 | 3.00 |

Explanation on the indexing serial number:

1--Baffle plates (may be below the bottom of the enclosure);

2---Bottom of the enclosure.

Y ≥ twice X but never less than 25 mm.

Figure 8- Baffle

Key:

A--Part or component of the equipment that is considered to be a source of fire HAZARD. This consists of an entire component or part of the equipment if it is not otherwise shielded, or the unshielded portion of a component that is partially shielded by its casing.

B--Projection of the outline of A on the horizontal plane.

C--Inclined line that traces out the minimum area of the bottom and sides to be constructed as specified in 9.3.2 c) 1) and 9.3.2 c) 2). This line projects at a 5° angle from the vertical at every point around the perimeter of A and is oriented so as to trace out the maximum area.

D--Minimum area of the bottom to be constructed as specified in 9.3.2 c) 1).

Figure 9 - Area of the bottom of an ENCLOSURE to be constructed as specified in 9.3.2 c) 1)

## 9.4	Limited-energy circuit

A limited-energy circuit is a circuit that meets all the following criteria.

a)	The potential appearing in the circuit is not more than 30 V (effective value) and 42.4 V (peak value), or 60 V d.c.

b)	The current that can appear in the circuit is limited by one of the following methods:

1)	the maximum available current is limited inherently or by impedance so that it cannot exceed the relevant specified value of Table 12;

2)	current is limited by an overcurrent protection device so that it cannot exceed the applicable values of Table 13;

3)	a regulating network limits the maximum available current so that it cannot exceed the relevant value of Table 12 in NORMAL CONDITION or as a result of a fault in the regulating network.

c)	It is separated by at least BASIC INSULATION from other circuits having energy values exceeding criteria a) or b) above.

If an overcurrent protection device is used, it shall be a fuse or a non-adjustable non-self-resetting electromechanical device.

Table 12 – Limits of maximum available current

| Open-circuit output voltage (U or Û) V |  |  | Maximum available current A |
| --- | --- | --- | --- |
| a.c. r.m.s. | Direct current | Peak a | a.c. r.m.s. or d.c.a.c. r.m.s. |
| U≤2 | U≤2 | Û ≤2.8 | 50 |
| 2<U≤12.5 | 2<U≤12.5 | 2.8< Û ≤17.6 | 100/U |
| 12.5<U≤18.7 | 12.5<U≤18.7 | 17.6< Û ≤26.4 | 8 |
| 18.7<U≤ 30 | 18.7<U≤60 | 26.4< Û ≤42.4 | 150/U |
| aThe peak value (Û) applies to non-sinusoidal a.c. and to d.c. with ripple exceeding 10 %, and is provided for convenience. The r.m.s. value of the maximum available current shall be determined as that value is related to heating. |  |  |  |

Table 13 - Values for overcurrent protection devices

| Potential appearing in the circuit (U or Û) V |  |  | Current at which the overcurrent protection device is disconnected after not more than 120sb,c A |
| --- | --- | --- | --- |
| a.c. r.m.s. | Direct current | Peak a | a.c. r.m.s. or d.c.a.c. r.m.s. |
| U≤2 | U≤2 | Û ≤2.8 | 62.5 |
| 2<U≤12.5 | 2<U≤12.5 | 2.8< Û ≤17.6 | 125/U |
| 12.5<U≤18.7 | 12.5<U≤18.7 | 17.6< Û ≤26.4 | 10 |
| 18.7<U≤30 | 18.7<U≤60 | 26.4< Û ≤42.4 | 200/U |
| a The peak value (Û) applies to non-sinusoidal a.c. and to d.c. with ripple exceeding 10 %, and is provided for convenience. The r.m.s. value of the maximum available current shall be determined as that value is related to heating. b The evaluation is based on the specified time-current breaking characteristic of the device, which is different from the RATED breaking current. (For example, an ANSI/UL 248-14 5 A fuse is specified to break 10 A at 120 s or less and an IEC 60127 T-type 4 A fuse is specified to break at 8,4 A at 120 s or less.) c The breaking current of fuses is dependent on temperature, and this has to be taken into account if the temperature immediately around the fuse is significantly higher than the room temperature. |  |  |  |

## 9.5	Requirements for equipment containing or using flammable liquids

Flammable liquids contained in, or specified for use with, the equipment shall not cause the spread of fire in NORMAL CONDITION or in SINGLE FAULT CONDITION.

The HAZARDS arising from flammable liquids are considered to be reduced to a tolerable level if one of the following requirements is met.

a)	The temperature of the surface of the liquid and parts in contact with the surface, in NORMAL CONDITION and SINGLE FAULT CONDITION, is limited to a temperature not exceeding t - 25 °C, where t is the fire point of the liquid.

Note:	Fire point is the temperature to which a liquid must be heated (under specified conditions) so that the vapour/air mixture at the surface will support a flame for at least 5 s when an external flame is applied and withdrawn.

b)	The quantity of liquid is limited to an amount that could not cause the spread of fire.

c)	If the liquid can be ignited, the flames are contained to prevent the spread of fire outside the equipment. Detailed instructions for use shall be provided to establish adequate RISK reduction procedures.

## 9.6	Overcurrent protection

### 9.6.1	General rule

Equipment intended to be energized from a MAlNS supply shall be protected by fuses, circuit­ breakers, thermal cut-outs, impedance limiting circuits or similar means, to provide protection against excessive current being drawn from the MAINS in case of a fault in the equipment.

A minimum of BASIC INSULATION is required between MAINS-connected parts of opposite polarity on the supply side of the overcurrent protection device.

Overcurrent protection devices shall not be fitted in the protective conductor. Fuses or single pole circuit-breakers shall not be fitted in the neutral conductor of multi-phase equipment.

### 9.6.2	PERMANENTLY CONNECTED EQUIPMENT

Overcurrent protection devices are optional in PERMANENTLY CONNECTED EQUIPMENT. If none are fitted, the installation instructions shall specify the characteristics of the overcurrent protection devices required in the building installation.

### 9.6.3	Other Equipment

If an overcurrent protection device is provided, it shall be within the equipment.

# 10	Equipment temperature limits and resistance to heat

## 10.1	Surface temperature limits for protection against burns

The temperature of easily touched surfaces shall not exceed the values of Table 14 in NORMAL CONDITION, and 105 °C in SINGLE FAULT CONDITION, at an ambient temperature of 40 °C.

Easily touched surfaces of equipment RATED for a maximum ambient temperature above 40 °C are permitted to exceed the values of Table 14 in NORMAL CONDITION, and to exceed 105 °C in SINGLE FAULT CONDITION, by not more than the amount by which the maximum RATED temperature exceeds 40 °C.

Surfaces protected by barriers that prevent them from being touched accidentally are not considered to be easily touched surfaces, provided that the barriers cannot be removed without the use of a TOOL.

Table 14 – Surface temperature limits in NORMAL CONDITION

| Part | Limit °C |
| --- | --- |
| 1. Outer surface of enclosure (unintentional contact) |  |
| a)Metal, uncoated or anodized | 65 |
| b) Metal, coated (painted, non-metallic) | 80 |
| c) Plastics | 85 |
| d) Glass and ceramics | 80 |
| e) Small areas (<2cm2 ) that are not likely to be touched in normal use | 100 |
| 2. Knobs and handles (to be touched in normal use) |  |
| a) Metal | 55 |
| b) Plastics | 70 |
| c) Glass and ceramics | 65 |
| d) Non-metallic parts that in NORMAL USE are held only for short periods (1 s - 4 s) | 70 |

## 10.2	Temperatures of windings

If a HAZARD could be caused by excessive temperature, the temperature of the insulating material of windings shall not exceed the values of Table 15 in NORMAL CONDITION and SINGLE FAULT CONDITION.

Table 15 - Maximum temperatures for insulation material of windings

| Class of insulation (see GB/T 11021) | NORMAL CONDITION °C | SINGLE FAULT CONDITION °C |
| --- | --- | --- |
| A | 105 | 150 |
| B | 130 | 175 |
| E | 120 | 165 |
| F | 155 | 190 |
| H | 180 | 210 |

## 10.3	Resistance to heat

CLEARANCES and CREEPAGE DISTANCES shall meet the requirements of 6.6 when the equipment is operated at an ambient temperature of 40 °C, or the maximum RATED ambient temperature if higher.

ENCLOSURE of non-metallic material shall be resistant to high temperatures. Insulating material shall have adequate resistance to heat.

# 11	Protection against HAZARDS from fluids and solid foreign objects

## 11.1	General rule

Equipment shall be designed to give protection to OPERATORS and the surrounding area against HAZARDS from fluids and solid foreign objects encountered in NORMAL USE.

Note: Fluids likely to be encountered fall into three categories:

a)	those having continuous contact, for example, in vessels intended to contain them;

b)	those having occasional contact, for example, cleaning fluids;

c)	those having accidental (unexpected) contact. The manufacturer cannot safeguard against such cases.

Fluids specified by the manufacturer, including cleaning and decontaminating fluids, are considered. Other fluids are not considered.

## 11.2	Cleaning

If a cleaning or decontamination process is specified by the manufacturer, this shall not cause a direct HAZARD, nor an electrical HAZARD, nor a HAZARD resulting from corrosion or other weakening of structural parts relied upon for safety.

## 11.3	Spillage

If, in NORMAL USE, liquid is likely to be spilt into the equipment, the equipment shall be designed so that no HAZARD will occur, for example as a result of the wetting of insulation or of internal uninsulated parts, or as a result of the contact of potentially aggressive substances (such as corrosive, toxic or flammable liquids) with parts of the equipment.

## 11.4	Overflow

Liquid overflowing from any container in the equipment which can be overfilled shall not cause a HAZARD during NORMAL USE, for example as a result of the wetting of insulation or of internal uninsulated parts that are HAZARDOUS LIVE.

Equipment likely to be moved while a container is full of liquid shall be protected against liquid surging out of the container.

## 11.5	Battery electrolyte

Batteries shall be so mounted that safety cannot be impaired by leakage of their electrolyte.

## 11.6	Equipment RATED with a degree of ingress protection (IP code)

Equipment RATED by the manufacturer as conforming to one of the degrees of protection as defined in GB/T 4208 shall adequately resist the ingress of solid foreign objects and water which could lead to a HAZARD.

## 11.7	Fluid pressure and leakage

### 11.7.1	Maximum Pressure

The maximum pressure to which a part of the equipment can be subjected in NORMAL USE or SINGLE FAULT CONDITION shall not exceed the RATED maximum working pressure for the part.

The maximum pressure shall be considered to be the highest of the following:

a)	the RATED maximum supply pressure specified for an external source;

b)	the pressure setting of an overpressure safety device provided as part of the assembly;

c)	the maximum pressure that can be developed by a pressure generating device that is part of the assembly, unless the pressure is limited by an overpressure safety device.

### 11.7.2	Leakage and rupture at high pressure

Fluid-containing parts which in NORMAL USE have both of the following characteristics shall not cause a HAZARD through rupture or leakage:

a)	The product of pressure and volume is greater than 200kPa·L.

b)	A pressure greater than 50 kPa.

### 11.7.3	Leakage from low-pressure parts

Leakage from fluid-containing parts at pressures lower than the levels of 11.7.2 shall notcause a hazard.

### 11.7.4	Overpressure safety device

An overpressure safety device shall not operate in NORMAL USE. It shall conform to the following requirements:

a)	It shall be connected as close as possible to the fluid-containing parts of the system that it is intended to protect;

b)	It shall be installed so as to provide easy access for inspection, maintenance and repair;

c)	It shall not be capable of being adjusted without the use of a TOOL;

d)	It shall have its discharge opening so located and directed that the released material is not directed towards any person;

e)	It shall have its discharge opening so located and directed that operation of the device will not deposit material on parts if that could cause a HAZARD;

f)	It shall have adequate discharge capacity to ensure that the pressure cannot exceed the RATED maximum working pressure of the system.

There shall be no shut-off valve between an overpressure safety device and the parts that it is intended to protect.

# 12	Protection against radiation (including laser source), sound pressure and ultrasonic pressure

## 12.1	General rule

The equipment shall provide protection against the effects of internally generated ultraviolet, ionizing and microwave radiation; laser sources, and sonic and ultrasonic pressure.

## 12.2	Equipment producing ionizing radiation

### 12.2.1	Ionizing radiation

#### 12.2.1.1	General rule

Equipment containing or generating ionizing radiation (from either radioactive sources or X­radiation) shall meet the following requirements.

a)	If it is intended to emit radiation, it shall meet the requirements of 12.2.1.2. Alternatively, if it is within the scope of GB/T 19661 (all parts), it shall be tested, classified, and marked according to GB/T 19661 (all parts).

b)	If it uses or generates radiation but only emits stray radiation, it shall meet the requirements of 12.2.1.3.

#### 12.2.1.2	Equipment intended to emit radiation

Equipment which contains radioactive substances or that generates x-rays, and is intended to emit ionizing radiation outside the equipment, if the effective dose rate at any easily reached point 50 mm from the outer surface exceeds 5 µSv/h, the equipment shall be marked with all the following markings:

a)	With symbol 16 of Table 1;

b)	For equipment containing one or more radioactive substances, with the abbreviations of the radionuclides;

c)	Either with the maximum dose rate value at 1 m, or with a dose rate value between 1 µSv/h and 5 µSv/h at the appropriate distance in m.

Example: “2,5 pSv/h at 1 m”; "3 pSv/h at 0,3 m"

#### 12.2.1.3	Equipment not intended to emit radiation

The effective dose rate of unintended stray radiation at any easily reached point 100 mm from the outer surface of the equipment shall not exceed 1 pSv/h.

### 12.2.2	Accelerated electrons

The equipment shall be so constructed that compartments in which electrons are accelerated by voltages exceeding 5 kV cannot be opened without the use of a TOOL.

## 12.3	Optical Radiations

Equipment with lamps and lamp systems emitting ultraviolet, visible, or infrared radiation, including light emitting diodes, shall not permit unintentional escape of radiation that could cause a HAZARD.

## 12.4	Microwave radiation

In normal condition and single fault condition, at any point 50 mm from the equipment, the power density of spurious microwave radiation at frequencies between 1 GHz and 100 GHz shall not exceed 10 W/m2. This requirement does not apply to parts of equipment intended to transmit microwave radiation, such as those located at the waveguide output ports.

## 12.5	Sonic and ultrasonic pressure

### 12.5.1	Sound pressure level

If equipment produces noise at a level which could cause a HAZARD, the manufacturer shall measure the maximum sound pressure level which the equipment can produce (except for sound from alarms and sound from parts remote from the equipment) and shall calculate the maximum sound power level.

Installation instructions shall specify how the RESPONSIBLE BODY can ensure that the sound pressure level from equipment, at its point of use after installation, will not reach a value which could cause a HAZARD. These instructions shall identify readily available and practicable protective materials or measures which may be used, including the fitting of noise-reducing baffles or hoods.

### 12.5.2	Ultrasonic pressure

If equipment not intended to emit ultrasound produces ultrasound at a level which could cause a HAZARD, the manufacturer shall measure the maximum ultrasonic pressure level which the equipment can produce.

If equipment intended to emit ultrasound produces ultrasonic pressure at a level which could cause a HAZARD, the manufacturer shall measure the maximum ultrasonic pressure level which the equipment can produce.

Outside the useful beam the ultrasonic pressure shall not exceed 110 dB above the reference pressure value of 20 µPa, for frequencies between 20 kHz and 100 kHz.

If inside the useful beam the ultrasonic pressure exceeds 110 dB for frequencies between 20 kHz and 100 kHz, the equipment shall be marked with symbol 14 of Table 1 and the following information shall be in the documentation:

a)	Dimensions of useful beam

b)	The area of the useful beam in which the ultrasonic pressure exceeds 110 dB;

c)	The maximum sound pressure value inside the beam area.

## 12.6	Laser sources

Equipment employing laser sources shall meet the requirements of GB 7247.1.

# 13	Protection against released gases and substances, explosions and implosions

## 13.1	Poisonous and injurious gases and substances

Equipment shall not liberate dangerous amounts of hazardous substances in NORMAL CONDITION and in SINGLE FAULT CONDITION.

If potentially-hazardous substances are liberated, the OPERATOR shall not be directly exposed to a quantity of the substance that could cause harm.

If NORMAL operation of the equipment requires the discharge of hazardous substances, and if that discharge is intended to be managed by the RESPONSIBLE BODY in accordance with the manufacturer's instructions, then such discharge is not considered to be liberation of hazardous substances.

## 13.2	Explosion and implosion

### 13.2.1	Component

If components liable to explode if overheated or overcharged are not provided with a pressure release device, protection for OPERATORS shall be incorporated in the equipment.

Pressure release devices shall be located so that a discharge will not cause danger to OPERATORS. The construction shall be such that no pressure release device is obstructed.

### 13.2.2	Batteries and battery charging

Batteries shall not cause explosion or produce a fire HAZARD as a result of excessive charge or discharge, or if a battery is installed with incorrect polarity. If necessary, protection shall be incorporated in the equipment, unless the manufacturer's instructions specify that it is for use only with batteries which have built-in protection.

If an explosion or fire HAZARD could occur through fitting a battery of the wrong type (for example, if a battery with built-in protection is specified) there shall be a warning marking on or near the battery compartment or mounting, and a warning in the manufacturer's instructions. An acceptable marking is symbol 13 of Table 1.

If the equipment has means for charging rechargeable batteries, and if non-rechargeable cells could be fitted and connected in the battery compartment, there shall be a warning marking in or near the compartment. The marking shall warn against the charging of non-rechargeable batteries and indicate the type of rechargeable battery that can be used with the recharging circuit. An acceptable marking is symbol 13 of Table 1.

The battery compartment shall be designed so that there is no possibility of explosion or fire caused by build-up of flammable gases.

### 13.2.3	Implosion of cathode ray tubes

Cathode-ray tubes with a maximum face dimension exceeding 160 mm shall be intrinsically protected against the effects of implosion and against mechanical impact, unless an ENCLOSURE provides adequate protection.

A non-intrinsically protected tube shall be provided with an effective protective screen which cannot be removed without the use of a TOOL. If a separate screen of glass is used, it shall not be in contact with the surface of the tube.

A cathode-ray tube is considered to be intrinsically protected with respect to the effects of implosion if no additional protection is necessary when it is correctly mounted.

# 14	Components and subassemblies

## 14.1	General rule

If safety is involved, components and subassemblies, such as power supplies and built-in information technology equipment, shall be used in accordance with their specified RATINGS unless a specific exception is made.

## 14.2	Motors

### 14.2.1	Motor temperatures

Motors which, when stopped or prevented from starting, would present an electric shock HAZARD, a temperature HAZARD, or a fire HAZARD shall be protected by an overtemperature or thermal protection device meeting the requirements of 14.3.

### 14.2.2	Series excitation motors

If an overspeeding series excitation motor could cause a HAZARD, it shall be connected directly to the devices driven by it.

## 14.3	Overtemperature protection devices

Overtemperature protection devices are devices operating in SINGLE FAULT CONDITION. They shall meet all of the following requirements:

a)	Be constructed so that reliable function is ensured;

b)	Be RATED to interrupt the maximum voltage and current of the circuit in which they are employed;

c)	Not operate in NORMAL USE.

If a self-resetting overtemperature protection device is used to prevent a HAZARD in case of failure of a temperature control system (for example a thermostat), the protected part of the equipment shall require intervention before becoming operational again.

## 14.4	Fuse holders

If a fuse is intended to be replaced by an OPERATOR the fuse holder shall not permit HAZARDOUS LIVE parts to be ACCESSIBLE during fuse replacement.

## 14.5	MAINS voltage selecting devices

Devices shall be constructed so that a change from one voltage or one type of supply to another cannot occur accidentally. See 5.1. 3d).

## 14.6	Printed wiring boards

Printed wiring boards shall be made of material with a flammability classification of V-1 or better of GB/T 5169.16.

## 14.7	Circuits used to limit TRANSIENT OVERVOLTAGES

Any overvoltage-limiting circuit constituting a part of the equipment shall have sufficient strength to limit possible transient overvoltage.

# 15	Protection by interlocks

## 15.1	General rule

Interlocks used to protect OPERATORS from HAZARDS shall prevent an OPERATOR from being exposed to the HAZARD before the HAZARD is removed and shall meet the requirements of 15.2 and 15.3.

## 15.2	Prevention of reactivating

An interlock system for the protection of OPERATORS shall ensure that a single fault in the interlock system is either unlikely to occur during the expected life of the equipment, or cannot cause a HAZARD.

## 15.3	Reliability

An interlock system for the protection of OPERATORS shall ensure that a single fault in the interlock system is either unlikely to occur during the expected life of the equipment, or cannot cause a HAZARD.

# 16	Test Method

The test methods corresponding to the technical requirements of Chapters 5 to 15 are shown in Table 16.

Table 16 Test methods

| No. | Subclauses | Test Method |
| --- | --- | --- |
| 1 | Chapter 5 Signs and documents | Carry out the test according to the test methods specified in Chapter 5 of GB/T42125.1-2024 |
| 2 | Chapter 6 Protection against electric shock | Carry out the test according to the test methods specified in Chapter 6 of GB/T42125.1-2024 |
| 3 | Chapter 7 Protection against mechanical hazards | Carry out the test according to the test methods specified in Chapter 7 of GB/T42125.1-2024 |
| 4 | Chapter 8 Resistance to mechanical stresses | Carry out the test according to the test methods specified in Chapter 8 of GB/T42125.1-2024 |
| 5 | Chapter 9 Protection against the spread of fire | Carry out the test according to the test methods specified in Chapter 9 of GB/T42125.1—2024 |

Table 16 Test methods (continued)

| No. | Subclauses | Test Method |
| --- | --- | --- |
| 6 | Chapter 10 Equipment temperature limits and resistance to heat | Carry out the test according to the test methods specified in Chapter 10 of GB/T42125.1—2024 |
| 7 | Chapter 11 Protection against HAZARDS from fluids and solid foreign objects | Carry out the test according to the test methods specified in Chapter 11 of GB/T42125.1—2024 |
| 8 | Chapter 12 Protection against radiation (including laser source), sound pressure and ultrasonic pressure | Carry out the test according to the test methods specified in Chapter 12 of GB/T42125.1—2024 |
| 9 | Chapter 13 PROTECTION AGAINST LIBERATED GASES AND SUBSTANCES, EXPLOSION AND IMPOSION | Carry out the test according to the test methods specified in Chapter 13 of GB/T42125.1—2024 |
| 10 | Chapter 14 Component and part | Carry out the test according to the test methods specified in Chapter 14 of GB/T42125.1—2024 |
| 11 | Chapter 15 Protection by interlocks | Carry out the test according to the test methods specified in Chapter 15 of GB/T42125.1—2024 |

# Appendix A 
(normative) 
Insulation requirements not covered by 6.6

## A.1	Insulation for mains circuits

A.1.1	General rule

If a manufacturer rates equipment for OVERVOLTAGE CATEGORY Ⅲ OR OVERVOLTAGE CATEGORY Ⅳ, the relevant requirements of this annex apply. OR OVERVOLTAGE CATEGORY Ⅳ, the relevant requirements of this annex apply.

OVERVOLTAGE CATEGORY Ⅱ is for equipment intended to be supplied from the building wiring. It applies both to plug-connected equipment and to PERMANENTLY CONNECTED EQUIPMENT. Subclause 6.6 covers only the requirements for OVERVOLTAGE CATEGORY Ⅱ with a nominal supply voltage up to 300 V. The requirements for higher OVERVOLTAGE CATEGORIES and for OVERVOLTAGE CATEGORY Ⅱ with a nominal supply voltage above 300 V are covered by this annex.

OVERVOLTAGE CATEGORY Ⅲ is for equipment intended to form part of a building wiring installation. Such equipment includes socket outlets, fuse panels, and some MAINS installation control equipment. Manufacturers may also design equipment for OVERVOLTAGE CATEGORY Ⅲ when an even higher degree of reliability and availability is desired.

OVERVOLTAGE CATEGORY Ⅳ is for equipment installed at or near the origin of the electrical supply to a building, between the building entrance and the main distribution board. Such equipment may include electricity tariff meters and primary overcurrent protection devices. Manufacturers may also design equipment for OVERVOLTAGE CATEGORY Ⅳ when an even higher degree of reliability and availability is desired.

A.1.2 	CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS

CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS shall meet the values of the following tables as applicable:

a)	for MAINS CIRCUITS of OVERVOLTAGE CATEGORY II, with nominal supply voltages above 300 V, Table A.2;

b)	for MAINS CIRCUITS of OVERVOLTAGE CATEGORY Ill, Table A.3;

c)	for MAINS CIRCUITS of OVERVOLTAGE CATEGORY IV, Table A.4.

The values in the following table A.2～table A.4 are for BASIC INSULATION and SUPPLEMENTARY INSULATION. Values for REINFORCED INSULATION shall be twice the values for BASIC INSULATION.

If the equipment is RATED to operate at an altitude greater than 2000 m, the CLEARANCES shall be multiplied by the applicable factor of Table A.1.

Table A.1 - Multiplication factors for CLEARANCES for equipment RATED for operation at altitudes up to 5 000 m

| RATED operating altitude m | Multiplication factor |
| --- | --- |
| ≤2000 | 1.00 |
| 2001~3000 | 1.14 |
| 3001~4000 | 1.29 |
| 4001~5000 | 1.48 |

Table A.2 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOL TAGE CATEGORY II above 300 V

| Voltage line-to-neutral a.c. r.m.s. or d.c. | CLEARANCES | CREEPAGE DISTANCES |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Printed wiring board material |  | Other insulating material |  |  |  |  |  |  |
|  |  | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 |  |  | POLLUTION DEGREE 3 |  |  |
|  |  | All material groups | Material groups Ⅰ, Ⅱ, Ⅲa | All material groups | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ |
| V | mm | mm | mm | mm | mm | mm | mm | mm | mm | mm |
| >300~600 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 4.3 | 6.0 | 7.5 | 8.3 | 9.4 |
| >600~1000 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 7.2 | 10.0 | 12.5 | 14.0 | 16.0 |

Table A.3 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOLTAGE CATEGORY Ⅲ

| Voltage line-to-neutral a.c. r.m.s. or d.c. | CLEARANCES | CREEPAGE DISTANCES |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Printed wiring board material |  | Other insulating material |  |  |  |  |  |  |
|  |  | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 |  |  | POLLUTION DEGREE 3 |  |  |
|  |  | All material groups | Material groups II, II, IIIa | All material groups | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ |
| V | mm | mm | mm | mm | mm | mm | mm | mm | mm | mm |
| ≤150 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 | 1.6 | 2.0 | 2.2 | 2.5 |
| >150~300 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.8 | 4.1 | 4.7 |
| >300~600 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 6.0 | 7.5 | 8.3 | 9.4 |
| >600~1000 | 8.0 | 8.0 | 8.0 | 8.0 | 8.0 | 8.0 | 10.0 | 12.5 | 14.0 | 16 |

Table A.4 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOLTAGE CATEGORY IV

| Voltage line-to-neutral a.c. r.m.s. or d.c. | CLEARANCES | CREEPAGE DISTANCES |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Printed wiring board material |  | Other insulating material |  |  |  |  |  |  |
|  |  | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 |  |  | POLLUTION DEGREE 3 |  |  |
|  |  | All material groups | Material groups II, II, IIIa | All material groups | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ |
| V | mm | mm | mm | mm | mm | mm | mm | mm | mm | mm |
| ≤150 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 | 3.0 |

Table A.4 - CLEARANCES and CREEPAGE DISTANCES for MAINS CIRCUITS of OVERVOLTAGE CATEGORY IV(continued)

| Voltage line-to-neutral a.c. r.m.s. or d.c. | CLEARANCES | CREEPAGE DISTANCES |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Printed wiring board material |  | Other insulating material |  |  |  |  |  |  |
|  |  | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 | POLLUTION DEGREE 1 | POLLUTION DEGREE 2 |  |  | POLLUTION DEGREE 3 |  |  |
|  |  | All material groups | Material groups Ⅰ, Ⅱ, Ⅲa | All material groups | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ |
| V | mm | mm | mm | mm | mm | mm | mm | mm | mm | mm |
| >150~300 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 | 5.5 |
| >300~600 | 8.0 | 8.0 | 8.0 | 8.0 | 8.0 | 8.0 | 8.0 | 8.0 | 8.3 | 9.4 |
| >600~1000 | 14.0 | 14.0 | 14.0 | 14.0 | 14.0 | 14.0 | 14.0 | 14.0 | 14.0 | 16.0 |

A.1.3	Solid insulation for MAINS CIRCUITS

A.1.3.1.	General rule

Solid insulation of MAINS CIRCUITS shall withstand the electrical and mechanical stresses that may occur in NORMAL USE, in all RATED environmental conditions (see Clause 4), during the intended life of the equipment.

Solid insulation shall also meet the following requirements, as applicable:

1)	for moulded and potted parts, the requirements of A.1.3.2;

2)	for inner insulating layers of printed wiring boards, the requirements of A.1.3.3;

3)	for thin-film insulation, the requirements of A.1.3.4;

A.1.3.2	Moulded and potted parts

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers moulded together (see Figure A.1, item L) shall be separated by at least the applicable minimum distance of Table A.5 after the moulding is completed.

Key:

1--Layer 1;

2--Layer 2;

C--Conductor;

L--Conductor spacing.

Figure A.1 - Distance between conductors on an interface between two layers

A.1.3.3	Inner insulating layers of printed wiring boards

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers (see Figure A.2, item L) shall be separated by at least the applicable minimum distance of Table A.5.

Key:

L--Distance between adjacent conductors;

A--Layer;

C--Conductor.

Figure A.2 - Distance between adjacent conductors along an interface of an inner layer

Table A.5 - Minimum values for distance or thickness of solid insulation

| Voltage line-to-neutral a.c. r.m.s. or d.c. V | Minimum thickness a mm | Minimum distance L (see Figure A.2) a, b mm |
| --- | --- | --- |
| ≤300 | 0.4 | 0.4 |
| >300~600 | 0.6 | 0.6 |
| >600~1000 | 1.0 | 1.0 |
| aThese values are independent of the overvoltage category. bThese values apply for BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION. |  |  |

REINFORCED INSULATION of inner insulating layers of printed wiring boards shall also have adequate electric strength through the respective layers. One of the following methods shall be used:

a)	the thickness through the insulation is at least the value of Table A.5;

b)	the insulation is assembled from at least two separate layers of printed wiring board materials, each of which is RATED by the manufacturer of the material for an electric strength of at least the value of the test voltage of the applicable Table A.6, Table A.7 or Table A.8 for BASIC INSULATION.

c)	the insulation is assembled from at least two separate layers of printed wiring board materials, and the combination of layers is RATED by the manufacturer of the material for an electric strength of at least the value of the applicable Table A.6, Table A.7 or Table A.8 for REINFORCED INSULATION.

Table A.6 - Test voltages for solid insulation in mains circuits of OVERVOLTAGE CATEGORY II above 300 V

| Line-to-neutral a.c. r.m.s. or d.c. | Test Voltage |  |  |  |
| --- | --- | --- | --- | --- |
|  | 5s AC test V |  | Impulse test V |  |
|  | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION |
| >300~600 | 2210 | 3510 | 4000 | 6400 |
| >600~1000 | 3310 | 5400 | 6000 | 9600 |

Table A.7 -Test voltages for solid insulation in MAINS CIRCUITS of OVERVOL TAGE CATEGORY Ⅲ

| Line-to-neutral a.c. r.m.s. or d.c. | Test Voltage |  |  |  |
| --- | --- | --- | --- | --- |
|  | 5s AC test V |  | Impulse test V |  |
|  | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION |
| ≤150 | 1390 | 2210 | 2500 | 4000 |
| >150~300 | 2210 | 3510 | 4000 | 6400 |
| >300~600 | 3310 | 5400 | 6000 | 9600 |
| >600~1000 | 4260 | 7400 | 8000 | 12800 |

Table A.8 -Test voltages for solid insulation in MAINS CIRCUITS of OVERVOL TAGE CATEGORY IV

| Line-to-neutral a.c. r.m.s. or d.c. | Test Voltage |  |  |  |
| --- | --- | --- | --- | --- |
|  | 5s AC test V |  | Impulse test V |  |
|  | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION | BASIC INSULATION and SUPPLEMENTARY INSULATION | REINFORCED INSULATION |
| ≤150 | 2210 | 3510 | 4000 | 6400 |
| >150~300 | 3310 | 5400 | 6000 | 9600 |
| >300~600 | 4260 | 7400 | 8000 | 12800 |
| >600~1000 | 6600 | 11940 | 12000 | 19200 |

A.1.3.4	Thin-film insulation

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers (see Figure A.3, item L) shall be separated by at least the applicable CLEARANCE and CREEPAGE DISTANCE of A.1.2.

Key:

L--Distance between adjacent conductors;

A--Layers of thin-film material, such as tape and polyester film;

C--Conductor.

There may be air present between the layers

Figure A.3 - Distance between adjacent conductors located between the same two layers

REINFORCED INSULATION through the layers of thin-film insulation shall also have adequate electric strength. One of the following methods shall be used.

a)	the thickness through the insulation is at least the value of Table A.5.

b)	the insulation is assembled from at least two separate layers of printed wiring board materials, and the combination of layers is RATED by the manufacturer of the material for an electric strength of at least the value of the applicable Table A.6, Table A.7 or Table A.8 for REINFORCED INSULATION.

c)	the insulation consists of at least three separate layers of thin-film materials, any two of which have been tested to exhibit adequate electric strength.

## A.2	Insulation in secondary circuits

A.2.1	General rule

In this standard, the isolation of the secondary circuits and the mains circuits is realized by a transformer, in which the primary windings are separated from the secondary windings by reinforced insulation, double insulation, or a screen connected to the protective terminal.

A.2.2	CLEARANCES

CLEARANCES for secondary circuits shall meet a) or b):

a)	for BASIC INSULATION or SUPPLEMENTARY INSULATION, meet the applicable values of Tables A.9, A.10 and A.11, or for REINFORCED INSULATION meet twice that value;

b)	pass the voltage test using the applicable value of Tables A.9, A.10 and A.11.

For the application of Tables A.9 to A.11, the following applies:

1)	values for test voltages for REINFORCED INSULATION are 1,6 times the values for BASIC INSULATION;

2)	if the equipment is RATED to operate at an altitude greater than 2 000 m, the values for CLEARANCES are multiplied by the applicable factor of Table A.1;

3)	minimum clearance is 0,2 mm for pollution degree 2 and 0,8 mm for pollution degree 3.

Table A.9 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOLTAGE CATEGORY II above 300 V

| Secondary circuit Working voltage |  | MAINS VOLTAGE >300~600V a.c. r.m.s. |  | MAINS VOLTAGE >600~1000V a.c. r.m.s. |  |
| --- | --- | --- | --- | --- | --- |
| a.c. r.m.s. | d.c. or a.c. peak | CLEARANCES | Test voltage V a.c. r.m.s. | CLEARANCES | Test voltage V a.c. r.m.s. |
| V | V | mm | V | mm | V |
| 16 | 22.6 | 1.5 | 1390 | 2.9 | 1590 |
| 33 | 46.7 | 1.5 | 1390 | 3.0 | 2210 |
| 50 | 70 | 1.5 | 1390 | 3.0 | 2210 |
| 100 | 140 | 1.6 | 1450 | 3.1 | 2260 |
| 150 | 210 | 1.6 | 1450 | 3.2 | 2300 |
| 300 | 420 | 1.8 | 1540 | 3.4 | 2400 |
| 600 | 840 | 2.4 | 1620 | 3.9 | 2630 |
| 1000 | 1400 | 3.5 | 2450 | 5.0 | 3110 |
| 1250 | 1750 | 4.2 | 2770 | 5.8 | 3430 |
| 1600 | 2240 | 5.2 | 3190 | 6.9 | 3850 |
| 2000 | 2800 | 6.5 | 3700 | 8.2 | 4330 |
| 2500 | 3500 | 8.1 | 4300 | 9.8 | 4920 |
| 3200 | 4480 | 10 | 4950 | 12 | 5780 |
| 4000 | 5600 | 12 | 5780 | 15 | 7000 |
| 5000 | 7000 | 16 | 7400 | 18 | 8200 |
| 6300 | 8820 | 20 | 8980 | 22 | 9700 |
| 8000 | 11200 | 26 | 11200 | 28 | 11900 |
| 10000 | 14000 | 33 | 13800 | 35 | 14500 |
| 12500 | 17500 | 42 | 16900 | 44 | 17600 |
| 16000 | 22400 | 55 | 21200 | 57 | 21900 |
| 20000 | 28000 | 71 | 26300 | 73 | 27000 |
| 25000 | 35000 | 91 | 32600 | 93 | 33200 |
| 32000 | 44800 | 120 | 41600 | 122 | 42200 |
| 40000 | 56000 | 154 | 52200 | 157 | 53100 |
| 50000 | 70000 | 199 | 66100 | 202 | 67000 |
| 63000 | 88200 | 261 | 85300 | 262 | 85600 |
| Linear interpolation is allowed |  |  |  |  |  |

Table A.10 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOL TAGE CATEGORY Ⅲ

| Secondary circuit Working voltage |  | MAINS VOLTAGE ≤150V a.c. r.m.s. |  | MAINS VOLTAGE >150~300V a.c. r.m.s. |  | MAINS VOLTAGE >300~600V a.c. r.m.s. |  | MAINS VOLTAGE >600~1000V a.c. r.m.s. |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| a.c. r.m.s. | d.c. or a.c. peak | CLEARANCES | Test voltage V a.c. r.m.s. | CLEARANCES | Test voltage V a.c. r.m.s. | CLEARANCES | Test voltage V a.c. r.m.s. | CLEARANCES | Test voltage V a.c. r.m.s. |
| V | V | mm | V | mm | V | mm | V | mm | V |
| 16 | 22.6 | 0.48 | 1100 | 1.5 | 1800 | 2.9 | 2820 | 5.4 | 4240 |
| 33 | 46.7 | 0.50 | 1100 | 1.5 | 1800 | 3.0 | 2900 | 5.4 | 4240 |
| 50 | 70 | 0.53 | 1120 | 1.5 | 1800 | 3.0 | 2900 | 5.5 | 4300 |
| 100 | 140 | 0.61 | 1170 | 1.6 | 1880 | 3.1 | 2960 | 5.6 | 4360 |
| 150 | 210 | 0.69 | 1200 | 1.6 | 1880 | 3.2 | 3020 | 5.7 | 4420 |
| 300 | 420 | 0.94 | 1360 | 1.8 | 2040 | 3.4 | 3140 | 6.0 | 4600 |
| 600 | 840 | 1.6 | 1880 | 2.4 | 2440 | 3.9 | 3440 | 6.6 | 4860 |
| 1000 | 1400 | 2.5 | 2500 | 3.5 | 3200 | 5.0 | 4000 | 7.4 | 5240 |
| 1250 | 1750 | 3.2 | 3020 | 4.2 | 3620 | 5.8 | 4480 | 8.1 | 5560 |
| 1600 | 2240 | 4.1 | 3560 | 5.2 | 4120 | 6.9 | 5040 | 9.3 | 6120 |
| 2000 | 2800 | 5.3 | 4180 | 6.5 | 4800 | 8.2 | 5620 | 11 | 7000 |
| 2500 | 3500 | 6.9 | 5040 | 8.1 | 5560 | 9.8 | 6320 | 12 | 7500 |
| 3200 | 4480 | 9.2 | 6080 | 10 | 6400 | 12 | 7500 | 15 | 9100 |
| 4000 | 5600 | 12 | 7500 | 12 | 7500 | 15 | 9100 | 17 | 10100 |
| 5000 | 7000 | 15 | 9100 | 16 | 9600 | 18 | 10600 | 20 | 11600 |
| 6300 | 8820 | 19 | 11200 | 20 | 11600 | 22 | 12600 | 25 | 14100 |
| 8000 | 11200 | 25 | 14100 | 26 | 14600 | 28 | 15500 | 31 | 16900 |
| 10000 | 14000 | 32 | 17400 | 33 | 17800 | 35 | 18700 | 38 | 20000 |
| 12500 | 17500 | 41 | 21500 | 42 | 21900 | 44 | 22800 | 47 | 24200 |
| 16000 | 22400 | 54 | 27200 | 55 | 27600 | 57 | 28400 | 60 | 29700 |
| 20000 | 28000 | 69 | 33500 | 71 | 34300 | 73 | 35200 | 76 | 36400 |
| 25000 | 35000 | 89 | 41600 | 91 | 42400 | 93 | 43200 | 96 | 44400 |
| 32000 | 44800 | 118 | 53000 | 120 | 53700 | 122 | 54500 | 125 | 55600 |
| 40000 | 56000 | 153 | 66100 | 154 | 66500 | 157 | 67600 | 160 | 68700 |
| 50000 | 70000 | 198 | 82400 | 199 | 82700 | 202 | 83800 | 205 | 84900 |
| 63000 | 88200 | 260 | 104000 | 261 | 104400 | 262 | 104700 | 265 | 105700 |
| Linear interpolation is allowed |  |  |  |  |  |  |  |  |  |

Table A.11 - CLEARANCES and test voltages for secondary circuits derived from MAINS CIRCUITS of OVERVOLTAGE CATEGORY Ⅳ

| Secondary circuit Working voltage |  | MAINS VOLTAGE ≤150V a.c. r.m.s. |  | MAINS VOLTAGE >150~300V a.c. r.m.s. |  | MAINS VOLTAGE >300~600V a.c. r.m.s. |  | MAINS VOLTAGE >600~1000V a.c. r.m.s. |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| a.c. r.m.s. | d.c. or a.c. peak | CLEARANCES | Test voltage V a.c. r.m.s. | CLEARANCES | Test voltage V a.c. r.m.s. | CLEARANCES | Test voltage V a.c. r.m.s. | CLEARANCES | Test voltage V a.c. r.m.s. |
| V | V | mm | V | mm | V | mm | V | mm | V |
| 16 | 22.6 | 1.5 | 1800 | 2.9 | 2820 | 5.4 | 4240 | 8.3 | 5680 |
| 33 | 46.7 | 1.5 | 1800 | 3.0 | 2900 | 5.4 | 4240 | 8.3 | 5680 |
| 50 | 70 | 1.5 | 1800 | 3.0 | 2900 | 5.5 | 4300 | 8.4 | 5740 |
| 100 | 140 | 1.6 | 1880 | 3.1 | 2960 | 5.6 | 4360 | 8.5 | 5800 |
| 150 | 210 | 1.6 | 1880 | 3.2 | 3020 | 5.7 | 4420 | 8.6 | 5860 |
| 300 | 420 | 1.8 | 2040 | 3.4 | 3140 | 6.0 | 4600 | 8.9 | 5960 |
| 600 | 840 | 2.4 | 2440 | 3.9 | 3440 | 6.6 | 4860 | 9.6 | 6240 |
| 1000 | 1400 | 3.5 | 3200 | 5.0 | 4000 | 7.4 | 5240 | 10 | 6400 |
| 1250 | 1750 | 4.2 | 3620 | 5.8 | 4480 | 8.1 | 5560 | 11 | 7000 |
| 1600 | 2240 | 5.2 | 4120 | 6.9 | 5040 | 9.3 | 6120 | 12 | 7500 |
| 2000 | 2800 | 6.5 | 4800 | 8.2 | 5620 | 11 | 7000 | 13 | 8100 |
| 2500 | 3500 | 8.1 | 5560 | 9.8 | 6320 | 12 | 7500 | 15 | 9100 |
| 3200 | 4480 | 10 | 6400 | 12 | 7500 | 15 | 9100 | 17 | 10100 |
| 4000 | 5600 | 12 | 7500 | 15 | 9100 | 17 | 10100 | 19 | 11200 |
| 5000 | 7000 | 16 | 9600 | 18 | 10600 | 20 | 11600 | 23 | 13100 |
| 6300 | 8820 | 20 | 11600 | 22 | 12600 | 25 | 14100 | 27 | 15100 |
| 8000 | 11200 | 26 | 14600 | 28 | 15500 | 31 | 16900 | 33 | 17800 |
| 10000 | 14000 | 33 | 17800 | 35 | 18700 | 38 | 20000 | 40 | 21000 |
| 12500 | 17500 | 42 | 21900 | 44 | 22800 | 47 | 24200 | 50 | 25500 |
| 16000 | 22400 | 55 | 27600 | 57 | 28400 | 60 | 29700 | 63 | 31000 |
| 20000 | 28000 | 71 | 34300 | 73 | 35200 | 76 | 36400 | 79 | 37600 |
| 25000 | 35000 | 91 | 42400 | 93 | 43200 | 96 | 44400 | 99 | 45400 |
| 32000 | 44800 | 120 | 53700 | 122 | 54500 | 125 | 55600 | 129 | 57100 |
| 40000 | 56000 | 154 | 66500 | 157 | 67600 | 160 | 68700 | 164 | 70100 |
| 50000 | 70000 | 199 | 82700 | 202 | 83800 | 205 | 84900 | 209 | 86300 |
| 63000 | 88200 | 261 | 104400 | 262 | 104700 | 265 | 105700 | 268 | 106800 |
| Linear interpolation is allowed |  |  |  |  |  |  |  |  |  |

A.2.3	CREEPAGE DISTANCES

Based on the working voltage applied to the insulated part, the creepage distance for basic insulation or additional insulation for secondary circuit shall meet the applicable values of Table A.12. Values for REINFORCED INSULATION are twice the values for BASIC INSULATION.

Table A.12 - CREEPAGE DISTANCES for secondary circuits

| Secondary working voltage AC RMS or DC | Printed wiring board material |  | Other insulating material |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Pollution Degree 1 | Pollution Degree 2 | Pollution Degree 1 | Pollution Degree 2 |  |  | Pollution Degree 3 |  |  |
|  | All material groups | Material groups I, II, or IIIa | All material groups | Material group Ⅰ | Material group Ⅱ | Material group Ⅲ | Material group Ⅰ | Material group Ⅱ | Material group IIIb |
| V | mm | mm | mm | mm | mm | mm | mm | mm | mm |
| 10 | 0.025 | 0.04 | 0.08 | 0.40 | 0.40 | 0.40 | 1.00 | 1.00 | 1.00 |
| 12.5 | 0.025 | 0.04 | 0.09 | 0.42 | 0.42 | 0.42 | 1.05 | 1.05 | 1.05 |
| 16 | 0.025 | 0.04 | 0.10 | 0.45 | 0.45 | 0.45 | 1.10 | 1.10 | 1.10 |
| 20 | 0.025 | 0.04 | 0.11 | 0.48 | 0.48 | 0.48 | 1.20 | 1.20 | 1.20 |
| 25 | 0.025 | 0.04 | 0.125 | 0.50 | 0.50 | 0.50 | 1.25 | 1.25 | 1.25 |
| 32 | 0.025 | 0.04 | 0.14 | 0.53 | 0.53 | 0.53 | 1.3 | 1.3 | 1.3 |
| 40 | 0.025 | 0.04 | 0.16 | 0.56 | 0.80 | 1.10 | 1.4 | 1.6 | 1.8 |
| 50 | 0.025 | 0.04 | 0.18 | 0.60 | 0.85 | 1.20 | 1.5 | 1.7 | 1.9 |
| 63 | 0.040 | 0.063 | 0.20 | 0.63 | 0.90 | 1.25 | 1.6 | 1.8 | 2.0 |
| 80 | 0.063 | 0.10 | 0.22 | 0.67 | 0.95 | 1.3 | 1.7 | 1.9 | 2.1 |
| 100 | 0.10 | 0.16 | 0.25 | 0.71 | 1.00 | 1.4 | 1.8 | 2.0 | 2.2 |
| 125 | 0.16 | 0.25 | 0.28 | 0.75 | 1.5 | 1.5 | 1.9 | 2.1 | 2.4 |
| 160 | 0.25 | 0.40 | 0.32 | 0.80 | 1.1 | 1.6 | 2.0 | 2.2 | 2.5 |
| 200 | 0.40 | 0.63 | 0.42 | 1.00 | 1.4 | 2.0 | 2.5 | 2.8 | 3.2 |
| 250 | 0.56 | 1.0 | 0.56 | 1.25 | 1.8 | 2.5 | 3.2 | 3.6 | 4.0 |
| 320 | 0.75 | 1.6 | 0.75 | 1.60 | 2.2 | 3.2 | 4.0 | 4.5 | 5.0 |
| 400 | 1.0 | 2.0 | 1.0 | 2.0 | 2.8 | 4.0 | 5.0 | 5.6 | 6.3 |
| 500 | 1.3 | 2.5 | 1.3 | 2.5 | 3.6 | 5.0 | 6.3 | 7.1 | 8.0 |
| 630 | 1.8 | 3.2 | 1.8 | 3.2 | 4.5 | 6.3 | 8.0 | 9.0 | 10.0 |
| 800 | 2.4 | 4.0 | 2.4 | 4.0 | 5.6 | 8.0 | 10.0 | 11 | 12.5 |
| 1000 | 3.2a | 5.0a | 3.2 | 5.0 | 7.1 | 10.0 | 12.5 | 14 | 16 |
| 1250 |  |  | 4.2 | 6.3 | 9.0 | 12.5 | 16 | 18 | 20 |
| 1600 |  |  | 5.6 | 8.0 | 11 | 16 | 20 | 22 | 25 |
| 2000 |  |  | 7.5 | 10.0 | 14 | 20 | 25 | 28 | 32 |
| 2500 |  |  | 10.0 | 12.5 | 18 | 25 | 32 | 36 | 40 |
| 3200 |  |  | 12.5 | 16 | 22 | 32 | 40 | 45 | 50 |
| 4000 |  |  | 16 | 20 | 28 | 40 | 50 | 56 | 63 |
| 5000 |  |  | 20 | 25 | 36 | 50 | 63 | 71 | 80 |
| 6300 |  |  | 25 | 32 | 45 | 63 | 80 | 90 | 100 |
| 8000 |  |  | 32 | 40 | 56 | 80 | 100 | 110 | 125 |
| 10000 |  |  | 40 | 50 | 71 | 100 | 125 | 140 | 160 |
| 12500 |  |  | 50 | 63 | 90 | 125 |  |  |  |
| 16000 |  |  | 63 | 80 | 110 | 160 |  |  |  |
| 20000 |  |  | 80 | 100 | 140 | 200 |  |  |  |
| 25000 |  |  | 100 | 125 | 180 | 250 |  |  |  |
| 32000 |  |  | 125 | 160 | 220 | 320 |  |  |  |
| 40000 |  |  | 160 | 200 | 280 | 400 |  |  |  |
| 50000 |  |  | 200 | 250 | 360 | 500 |  |  |  |
| 63000 |  |  | 250 | 320 | 450 | 600 |  |  |  |
| Linear interpolation is allowed |  |  |  |  |  |  |  |  |  |
| aFor voltages above 1 000 V, CREEPAGE DISTANCES on printed wiring board material are the same as for other insulators of the same Material Group. bMaterial group Ⅲb is not recommended for application in POLLUTION DEGREE 3 above 630 V. |  |  |  |  |  |  |  |  |  |

A.2.4	Solid insulation

A.2.4.1	General rule

Solid insulation in secondary circuits shall withstand the electrical and mechanical stresses that may occur in NORMAL USE, in all RATED environmental conditions (see Clause 4), during the intended life of the equipment.

Conformity is checked by both of the following tests:

a)	for solid insulation used as an ENCLOSURE or BARRIER, the requirements of Clause 8;

b)	for moulded and potted parts, the requirements of A.2.4.2;

c)	for inner insulating layers of printed wiring boards, the requirements of A.2.4.3;

d)	for thin-film insulation, the requirements of A.2.4.4.

A.2.4.2	Moulded and potted parts

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, conductors located between the same two layers shall be separated by the applicable minimum distance of Table A.13 (see Figure A.1, item L).

Table A.13 - Minimum values for distance or thickness (see A.2.4.2 to A.2.4.4)

| AC peak voltage or DC Operating voltage or repeated peak voltage kV | Minimum value mm | Peak value of the a.c. or d.c. WORKING VOLTAGE or recurring peak voltage kV | Minimum value mm |
| --- | --- | --- | --- |
| >0.0467~0.33 | 0.05 | >8.0~10 | 3.5 |
| >0.33~0.8 | 0.1 | >10~12 | 4.5 |
| >0.8~1.0 | 0.15 | >12~15 | 5.5 |
| >1.0~1.2 | 0.2 | >15~20 | 8 |
| >1.2~1.5 | 0.3 | >20~25 | 10 |
| >1.5~2.0 | 0.45 | >25~30 | 12.5 |
| >2.0~2.5 | 0.6 | >30~40 | 17 |
| >2.5~3.0 | 0.8 | >40~50 | 22 |
| >3.0~4.0 | 1.2 | >50~60 | 27 |
| >4.0~5.0 | 1.5 | >60~80 | 35 |
| >5.0~6.0 | 2 | >80~100 | 45 |
| >6.0~8.0 | 3 |  |  |

A.2.4.3	Inner insulating layers of printed wiring boards

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION conductors located between the same two layers shall be separated by the applicable minimum distance of Table A.13 (see Figure A.2, item L).

REINFORCED INSULATION of inner insulating layers of printed wiring boards shall also have adequate electric strength through the respective layers. One of the following methods shall be used.

a)	the thickness of the insulation is at least the applicable minimum distance of Table A.13.

b)	the insulation is assembled from at least two separate layers of printed wiring board materials, each of which is RATED by the manufacturer of the material for an electric strength at least the value of the test voltage of the applicable Table A.9 to A.11 for BASIC INSULATION.

c)	the insulation is assembled from at least two separate layers of printed wiring board materials, and the combination of layers is RATED by the manufacturer of the material for an electric strength at least the value of the test voltage of the applicable Table A.9, Table A.10 or Table A.11 multiplied by 1.6 for REINFORCED INSULATION.

A.2.4.4 	Thin-film insulation

For BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION conductors located between the same two layers (see Figure A.3, item L) shall be separated by at least the applicable CLEARANCE and CREEPAGE DISTANCE of A.2.2 and A.2.3.

REINFORCED INSULATION through the layers of thin-film insulation shall also have adequate electric strength. One of the following methods can be used.

a)	the thickness through the insulation is at least the applicable value of Table A.13.

b)	the insulation consists of at least two separate layers of thin-film materials, each of which is RATED by the manufacturer of the material for an electric strength at least the value of the test voltage of the applicable Table A.9, Table A.10 or Table A.11 for BASIC INSULATION.

c)	the insulation consists of at least three separate layers of thin-film materials, any two of which have been tested to exhibit adequate electric strength.

## A.3	Insulation in circuits not addressed in 6.6, Clause A.1 or Clause A.2

A.3.1	General rule

These circuits have one or more of the following characteristics:

a)	the maximum possible TRANSIENT OVERVOLTAGE is limited by the supply source or within the equipment (see A.4.) to a known level below the level assumed for the MAINS CIRCUIT;

b)	the maximum possible TRANSIENT OVERVOLTAGE is above the level assumed for the MAINS CIRCUIT;

c)	the WORKING VOLTAGE is the sum of voltages from more than one circuit, or is a mixed voltage;

d)	the WORKING VOLTAGE includes a recurring peak voltage that may include a periodic non-sinusoidal waveform or a non-periodic waveform that occurs with some regularity;

e)	the WORKING VOLTAGE has a frequency above 30 kHz.

In cases a) to c), CLEARANCES for BASIC INSULATION and SUPPLEMENTARY INSULATION are determined according to A.3.2.

In cases d) and e) CLEARANCES are determined according to A.3.3.

In all cases A.3.4 addresses CREEPAGE DISTANCE and A.3.5 solid insulation.

A.3.2 	CLEARANCE calculation

CLEARANCES for BASIC INSULATION and SUPPLEMENTARY INSULATION are determined from the following formula (A.1):

Electrical clearance = D1 + F × (D2-D1)	(A.1)

Where,

F	--the coefficient is determined according to one of formulas (A.2) and (A.3);

D 1 and D 2 are values taken from Table A.14 for U m,

Among them:

D 1 represents the CLEARANCE that would be applicable to a TRANSIENT OVERVOLTAGE with the shape of a 1.2 × 50 µs impulse;

D2 represents the clearance that would be applicable to the peak working voltage without any transient overvoltage; the electrical clearance of the reinforced insulation is 2 times the electrical clearance value of the basic insulation.

F = (1.25 ×U w /U m ) – 0.25	if U w /U m ≤0.2

…	(A.2 )

F =0	if U w /U m ≤0.2	(A.3)

In formula (A.2) and formula (A.3):

Um =Uw +Ut;

U w = the maximum peak value of the WORKING VOLTAGE;

U t = the maximum additional TRANSIENT OVERVOLTAGE.

If the equipment is RATED to operate at an altitude greater than 2 000 m, the CLEARANCES shall be multiplied by the applicable factor of Table A.1.

Minimum CLEARANCE, for BASIC INSULATION, SUPPLEMENTARY INSULATION and REINFORCED INSULATION, is 0.2 mm for POLLUTION DEGREE 2 and 0.8 mm for POLLUTION DEGREE 3.

Table A.14 - CLEARANCE values for the calculation of A.3.2

| Maximum voltage Um | CLEARANCES |  | Maximum voltage Um | CLEARANCES |  |
| --- | --- | --- | --- | --- | --- |
|  | D1 | D2 |  | D1 | D2 |
| V | mm | mm | V | mm | mm |
| 14.1~266 | 0.010 | 0.010 | 4000 | 2.93 | 6.05 |
| 283 | 0.010 | 0.013 | 4530 | 3.53 | 7.29 |
| 330 | 0.010 | 0.020 | 5660 | 4.92 | 10.1 |
| 354 | 0.013 | 0.025 | 6000 | 5.37 | 10.8 |
| 453 | 0.027 | 0.052 | 7070 | 6.86 | 13.1 |
| 500 | 0.036 | 0.071 | 8000 | 8.25 | 15.2 |
| 566 | 0.052 | 0.10 | 8910 | 9.69 | 17.2 |
| 707 | 0.081 | 0.20 | 11300 | 12.9 | 22.8 |
| 800 | 0.099 | 0.29 | 14100 | 16.7 | 29.5 |
| 891 | 0.12 | 0.41 | 17700 | 21.8 | 38.5 |
| 1130 | 0.19 | 0.83 | 22600 | 29.0 | 51.2 |
| 1410 | 0.38 | 1.27 | 28300 | 37.8 | 66.7 |
| 1500 | 0.45 | 1.40 | 35400 | 49.1 | 86.7 |
| 1770 | 0.75 | 1.79 | 45300 | 65.5 | 116 |
| 2260 | 1.25 | 2.58 | 56600 | 85.0 | 150 |
| 2500 | 1.45 | 3.00 | 70700 | 110 | 195 |
| 2830 | 1.74 | 3.61 | 89100 | 145 | 255 |
| 3540 | 2.44 | 5.04 | 100000 | 165 | 290 |
| Linear interpolation is allowed |  |  |  |  |  |

A.3.3	CLEARANCES in circuits having recurring peak voltages, or having working VOLTAGES with frequencies above 30 kHz

Clearances for basic insulation and supplementary insulation for circuits having recurring peak voltages but not subjected to frequencies above 30 kHz shall meet the values of the second column of Table A.15, using the recurring peak voltage as the index (see Figure A.4 for an example of repeated peak voltage).

Key:

A--peak value of repeated voltage;

B--Operating voltage value.

Figure A.4 - Example of recurring peak voltage

Clearances for basic insulation and supplementary insulation for circuits that are subjected to frequencies above 30 kHz shall meet the values of the third column of Table Clearances for basic insulation and supplementary insulation for circuits that are subjected to frequencies above 30 kHz shall meet the values of the third column of Table A.15, using the peak value of the working voltage as the index.

Clearances for basic insulation and supplementary insulation for circuits that may be subjected to both recurring peak voltages and to frequencies above 30 kHz shall meet the higher of these requirements.

Clearances for reinforced insulation are twice the values for basic insulation.

If the equipment is RATED to operate at an altitude greater than 2000 m, the CLEARANCES shall be multiplied by the applicable factor of Table A.1.

The minimum clearance, for basic insulation, supplementary insulation and reinforced insulation, is 0.2 mm for pollution degree 2 and 0.8 mm for pollution degree 3.

Table A.15 - Clearances for basic insulation in circuits having recurring peak voltages or working voltages with frequencies above 30 kHz

| Voltage peak value | CLEARANCES |  |
| --- | --- | --- |
|  | Frequencies up to 30 kHz | Frequency above 30kHz |
| V | mm | mm |
| 0~330 | 0.01 | 0.02 |
| 400 | 0.02 | 0.04 |
| 500 | 0.04 | 0.07 |
| 600 | 0.06 | 0.11 |
| 800 | 0.13 | 0.26 |
| 1000 | 0.26 | 0.48 |
| 1200 | 0.42 | 0.76 |
| 1500 | 0.76 | 1.1 |
| 2000 | 1.27 | 1.8 |
| 2500 | 1.8 | 2.6 |
| 3000 | 2.4 | 3.5 |

Table A.15 - Clearances for basic insulation in circuits having recurring peak voltages or working voltages with frequencies above 30 kHz(continued)

| Voltage peak value | CLEARANCES |  |
| --- | --- | --- |
|  | Frequencies up to 30 kHz | Frequency above 30kHz |
| V | mm | mm |
| 4000 | 3.8 | 5.7 |
| 5000 | 5.7 | 8 |
| 6000 | 7.9 | 10 |
| 8000 | 11 | 15 |
| 10000 | 15.2 | 20 |
| 12000 | 19 | 25 |
| 15000 | 25 | 32 |
| 20000 | 34 | 44 |
| 25000 | 44 | 58 |
| 30000 | 55 | 72 |
| 40000 | 77 | 100 |
| 50000 | 100 |  |
| Linear interpolation is allowed |  |  |

A.3.4	CREEPAGE DISTANCES

The requirements of A.2.3 apply.

A.3.5	Solid insulation

The requirements of A.2.4 apply except that the values of Table A.16 are used in place of the applicable values of Table A.9, Table A.10, or Table A.11.

To determine the required test voltage from Table A.16, the following procedure shall be applied.

a)	Considering the requirements of A.3.3, calculate the theoretically required clearance according to A.3.2. Minimum electrical clearance is not applicable to contamination levels 2 and 3.

b)	Application of the resulting theoretically required CLEARANCE value of Table A.16 to determine the required test voltage.

Table A.16 Test voltages based on clearances

| Required Clearance mm | Test Voltage |  |
| --- | --- | --- |
|  | Impulse 1.2/50 ps | a.c. r.m.s. 50/60 Hz V |
| 0.010 0.025 0.040 0.063 0.1 0.2 0.3 0.5 1.0 1.5 2.0 2.5 3.0 4.0 4.5 5.0 6.0 8.0 10.0 12.0 15.0 20 25 30 40 50 60 80 100 | 330 440 520 600 810 1150 1310 1550 1950 2560 3090 3600 4070 4930 5330 5720 6460 7840 9100 10600 12900 16400 19900 23300 29800 36000 42000 53700 65000 | 230 310 370 420 500 620 710 840 1060 1390 1680 1960 2210 2680 2900 3110 3510 4260 4950 5780 7000 8980 10800 12700 16200 19600 22800 29200 35400 |
| Linear interpolation is allowed |  |  |

# REFERENCES

[1]	GB/T 2900.73-2008 Electrotechnical terminology - Earthing and protection against electric shock

[2]	GB/T 11021 Electrical insulation―Thermal classification

[3]	GB/T 5465.2 Graphical symbols for use on electrical equipment - Part 2: Graphical symbols

[4]	ISO 361 Basic ionizing radiation symbol

[5]	ISO 7000 Graphical symbols for use on equipment—Registered symbols

[6]	ISO 13854 Safety of machinery—Minimum gaps to avoid crushing of parts of the humanbody

[7]	ISO 13857 Safety of machinery—Safety distances to prevent hazard zones being reached by upper and lower limbs

[8]	IEC 60127(all parts) Miniature fuses

[9]	ANSI/UL 248-14 Low-Voltage Fuses—Part 14: Supplemental Fuses

[10]	EN 349 Safety of machinery .Minimum gaps to avoid crushing of parts of the human body

_________________

No part of this publication can be reproduced or transmitted in any form without the prior written permission of the publisher. Beijing Standard Translation Co., Ltd. reserves all the rights for final interpretation.